TENDER DOCUMENT

GOVERNMENT OF MADHYA PRADESH

DEPARTMENT OF MEDICAL EDUCATION

TENDER NO.			:	CPC/EQP-MOT/13-14/03

FOR ESTABLISHMENT OF :	MODULAR OPERATION THEATRE

AT GOVERNMENT / AUTONOMOUS MEDICAL COLLEGES, DENTAL COLLEGE & ASSOCIATED HOSPITALS UNDER MEDICAL EDUCATION DEPARTMENT, GOVT. OF M.P. BHOPAL

CONTRACT DOCUMENT CONSISTING OF TENDER NOTICE, TENDER FORM, RATE SHEET, TENDER CONDITIONS, SPECIFICATIONS AND TECHNICAL PARTICULARS, FORM OF AGREEMENT ETC.

ISSUING AUTHORITY

CHAIRMAN, PURCHASE COMMITTEE &

DEAN, M.G.M MEDICAL COLLEGE,

A.B. ROAD INDORE (M.P.)

INDEX

| Sr. | Particular | Page No. |
| --- | --- | --- |
|  | Important dates | 3 |
|  | Guideline to bidders for implementation of e-procurement system | 4 |
|  | Section- I (Invitation for bid) | 9 |
|  | Section –II (Instructions to bidder) | 11 |
|  | Section –III (General Conditions of Contract) | 29 |
|  | Section –IV(Special Conditions of Contract) | 45 |
|  | Section –V (Bid Form) | 54 |
|  | Table of Annexure | 55 |
|  | Requirement of equipments | 73 |

GOVERNMENT OF MADHYA PRADESH

DEPARTMENT OF MEDICAL EDUCATION

NATIONAL COMPETITIVE BIDDING FOR ESTABLISHMENT OF MODULAR OPERATION THEATRE AT

GOVERNMENT / AUTONOMOUS MEDICAL COLLEGES, DENTAL COLLEGE & ASSOCIATED HOSPITALS UNDER MEDICAL EDUCATION DEPARTMENT, BHOPAL, M.P. UNDER RATE CONTRACT

BID REFERENCE				:     	CPC/EQP-MOT/13-14/03

PRE BID MEETING				:	29.04.2013 AT 11:30 AM

Key Dates : Single Submission MultiOpening With Prequalification

| Seq No | DME Stage | Contractor Stage | Start Date & Time | Expiry Date & Time |
| --- | --- | --- | --- | --- |
| 1 | Tender Preparation and Release of NIT | - | 16-04-2013 10:00 | 17-04-2013 20:30 |
| 2 | - | Tender Purchase - Online | 17-04-2013 20:31 | 24-05-2013 17:30 |
| 3 | - | Tender Download | 17-04-2013 20:31 | 27-05-2013 17:00 |
| 4 | - | Submit Bid - Hash Online | 17-04-2013 20:31 | 29-05-2013 17:00 |
| 5 | Close for Bidding - Generation of Super Hash | - | 31-05-2013 10:00 | 31-05-2013 17:00 |
| 6 | - | Submit Bids Online - Re-encryption | 31-05-2013 17:31 | 01-06-2013 17:00 |
| 7 | Open EMD & Technical / PQ bid | - | 04-06-2013 10:00 | 04-06-2013 17:00 |
| 8 | Open Financial / Price-Bid | - | 12-06-2013 10:00 | 19-06-2013 17:00 |

| Last date for submission of original EMD and other relevant documents is 03.06.2013 upto 17:30 pm. All key dates are tentative and subject to be changed as per procurer requirement, all the information regarding this will be floated only on http://dme.mpeprocurement.gov.in portal. |  |
| --- | --- |

PLACE OF OPENING OF BIDS	 : Tender will be open online as per key dates

at http://dme.mpeprocurement.gov.in

ADDRESS FOR SUBMISSION OF 	 : Chairman, Purchase Committee & Dean

PHYSICAL TENDER  DOCUMENTS 	    M.G.M. Medical College,

A.B. Road, Indore – 452 001 (M.P.)

Telephone: 0731-2438271, 2526112,

2527383

Fax : 0731-2514628

Website: http://dme.mpeprocurement.gov.in ,

Email: cpc2013mgmmc.indore@yahoo.com,

Note: Tender can be downloaded from the website “http://dme.mpeprocurement.gov.in by making on line payment of Rs. 610/- (Rs. Six Hundred Ten only)

and

Cost of tender form in the form of  DD of Rs.5,000/-  in the name of Dean, M.G.M. Medical College, Indore payable at Indore (M.P.) should be submitted in a separate Envelope marked “Tender Fee” along with physical documents, failing which the tender will be rejected. Email: cpc2013mgmmc.indore@yahoo.com

GUIDELINES TO BIDDERS FOR IMPLEMENTATION OF

e-PROCUREMENT SYSTEM IN DEPARTMENT OF MEDICAL EDUCATION, GOVERNMENT OF MADHYA PRADESH

http://dme.mpeprocurement.gov.in

Note: These conditions will over-rule the conditions stated in the tender documents, wherever relevant and applicable.

- Registration of Bidders on e-Procurement System:

All the Bidders (Contractors/Vendors) registered / intending to register with DEPARTMENT OF MEDICAL EDUCATION are required to register on the e-Procurement System on the website http://mpeprocurement.gov.in and get empanelled on http://dme.mpeprocurement.gov.in in order to participate in restricted tenders processed by DEPARTMENT OF MEDICAL EDUCATION using the e-Procurement System.

The Bidders (Contractors) registered with other departments who are also eligible to participate in tenders processed by DEPARTMENT OF MEDICAL EDUCATION are also required to be registered on the e-Procurement System on http://mpeprocurement.gov.in.

Bidders who are already registered on http://mpeprocurement.gov.in can use their existing User Id to login and participate in online tenders floated by DEPARTMENT OF MEDICAL EDUCATION

Digital Certificates:

The Bids submitted online should be signed electronically with a Class III Digital Certificate to establish the identity of the Bidder submitting the Bid online. The Bidders may obtain Class III Digital Certificates issued by an approved Certifying Authority authorized by the Controller of Certifying Authorities, Government of India.

A Class III Digital Certificate is issued upon receipt of mandatory identity proofs along with an Application Form. Only upon the receipt of the required documents, a Digital Certificate can be issued.

Note: It may take up to 4 working days for issuance of Class III Digital Certificate, hence the bidders are advised to obtain them at the earliest.

The Bidders may obtain more information on Digital Certificate from

NexTenders (India) Pvt. Ltd.

A - 6, Sahyog Parisar Opp. Makhanlal Chaturvedi University,

E - 8, Shahpura,Bhopal - 462 039 (M.P.) Ph. No.: 0755 - 407 50 10

Email: Bhopal@nextenders.com

Important Note:

Bid for a particular tender can be submitted during the ONLINE BID SUBMISSION stage only using the Digital Certificate that is used to encrypt the data and sign the hash during the ONLINE BID PREPARATION AND HASH SUBMISSION stage.

In case, during the process of a particular tender, the bidder looses his Digital Certificate because of any problem (such as misplacement, virus attack, hardware problem, operating system problem, etc.); he may not be able to submit his bid online.

Hence, the bidders are advised to keep their Digital Certificates in a safe place under proper security to be used whenever required.

The digital certificate issued to the Authorized User of a Partnership firm / Private Limited Company / Public Limited Company / Joint Venture and used for online bidding will be considered as equivalent to a no-objection certificate / power of attorney to that user.

In case of Partnership Firm, majority of the Partners have to authorize a specific individual through Authority Letter signed by majority of the Partners of the firm.

In case of Private Limited Company, Public Limited Company, the Managing Director / any other Person (having designated authority to authorize a specific individual) has to authorize a specific individual through Authority Letter.

Unless the certificate is revoked, it will be assumed to represent adequate authority of the specific individual to bid on behalf of the Organization for online tenders as per Information Technology Act 2000. This Authorized User will be required to obtain a Digital Certificate. The Digital Signature executed through the use of Digital Certificate of this Authorized User will be binding on the firm. It shall be the responsibility of Management / Partners of the concerned firm to inform the Certifying Authority, if the Authorized User changes, and apply for a fresh digital certificate for the new Authorized User.

- Set Up of Bidders’ Computer System:

In order for a Bidder to operate on the e-Procurement System, the Computer System of the Bidder is required to be set up for Operating System, Internet Connectivity, Utilities, Fonts, etc. A help file on setting up of the Computer System can be obtained from e-Procurement Cell.

- Publishing of N.I.T.: For the Tenders processed using the e-Procurement System, only a brief Advertisement related to the Tender shall be published in the newspapers and the Detailed Notice shall be published only on the e-Procurement System. The contractors can view the Detailed Notice and the time schedule for all the Tenders processed using the e-Procurement System on the website http://dme.mpeprocurement.gov.in

- Key Dates: The Bidders are strictly advised to follow the time schedule (Key Dates) of the Tender for their side for tasks and responsibilities to participate in the Tender, as all the stages of each Tender are locked before the start time and date and after the end time and date for the relevant stage of the Tender as set by the concerned Department Official. However, it is advised to complete the required stage well in advance and not to wait for last date and time for processing.

- Online Processing Fee:

The Tender documents can be downloaded ONLY online. The Tender shall be available for download to the concerned eligible bidders immediately after online release of the Tender and upto the scheduled date and time as set in the key dates.

The Bidders can submit the tender documents online by making online payment of online processing fee of Rs.610/- using the service of the secure electronic payments gateway (see point 6 below for further details), and should print out the system generated receipt for their reference which can be produced whenever required. Arrangements have been made for the Bidders to make payments online (Rs.610/-) via Debit Card / Internet Banking towards online processing fee.

- Electronic Payment Account:

As the tender documents shall be available for download only online, Bidders are required to pay the online processing fee online using the Online Payments Gateway Service integrated into the e-Procurement System.

For the list of available modes of electronic payments that are presently accepted on the Online Payments Gateway Service, please check the link ‘List of e-Payments accepted Online’ on http://mpeprocurement.gov.in.

- Preparation of Bids and Submission of Bid Seals (Hashes) of Online Bids by the Bidders:

The Bidders have to prepare their Bids online, encrypt their Bid Data in the Bid Forms and submit Bid Seals (Hashes) of all the Envelopes and documents required to be uploaded related to the Tender as per the time schedule mentioned in the key dates of the Notice Inviting Tenders after signing of the same by the Digital Signature of their authorized representative.

- Generation of Super - Hash:

After the time of submission of Bid Seal (Hash) by the Contractors has lapsed, the bid round will be closed and the concerned Department Official shall generate a Tender Super – Hash which shall be digitally signed by the Department Official.

- Decryption and Re - Encryption Bids:

Bidders have to decrypt the Bid Data with their Digital Certificate and re-encrypt the Bid Data and also upload the relevant documents using Online Briefcase feature for which they had generated the Bid Seals (Hashes) during the Bid Preparation and Hash Submission stage within the specified time as stated in the time schedule (Key Dates).

The encrypted data of only those Bidders who have submitted their Bid Seals (Hashes) within the stipulated time, as per the tender time schedule (Key Dates), will be accepted by the system. A Bidder who has not submitted Bid Seals (Hashes) within the stipulated time will not be allowed to decrypt and re-encrypt his Bid Data or upload the relevant documents.

- Submission of Earnest Money Deposit:

The Bidders shall submit their Earnest Money Deposit as usual in a physically sealed Earnest Money Deposit Envelope and the same should reach the concerned Office as stated in the Notice Inviting Tender by post.

The Bidders have to upload scanned copy of Earnest Money Deposit instrument along with the reference details.

- Opening of Tenders:

The concerned Department Official receiving the tenders or his duly authorized Officer shall first open the online Earnest Money Deposit Envelope of all the Bidders and verify the scanned copy of the Earnest Money Deposit uploaded by the Bidders. He/She shall check for the validity of Earnest Money Deposit as required. He/She shall also verify the scanned documents uploaded by the Bidders, if any, as required. In case, the requirements are incomplete, the commercial and other Envelopes of the concerned Bidders received online shall not be opened.

The concerned Official shall then open the other Envelopes submitted online by the Bidders in the presence of the Bidders or their authorized representatives who choose to be present in the order of opening.

He/She will match the Bid Seal (Hash) of each Envelope and the documents uploaded, during the respective opening, with the hash submitted by the Bidders during the Bid Preparation and Hash Submission stage. In the event of a mismatch, the Bid Data in question shall be liable for a due process of verification by the Nodal Officer of e-Procurement System of DEPARTMENT OF MEDICAL EDUCATION.

SECTION I :

INVITATION FOR BIDS (IFB)

SECTION I : INVITATION FOR BIDS (IFB)

Date  :17.04.2013		                                         	  	IFB No.: CPC/EQP-MOT/13-14/03

- For the requirement of Indore, Bhopal, Jabalpur, Gwalior, Rewa & Sagar Medical Colleges, Dental College & Associated Hospitals Under Medical Education Department, Government of Madhya Pradesh, Bhopal Chairman, Purchase Committee & Dean, M.G.M. Medical College Indore (M.P.)  now invite online bids from eligible bidders for supply of Equipments For Modular Operation Theatre  under Rate Contract.

- Bids may be submitted by the primary manufacturer or their authorized distributor or importer for and on behalf of the primary manufacturer provided the bid is accompanied by a duly notarized letter of authority from the primary manufacturer.

  - Price of bidding document

(Non refundable DD)				:	Rs. 5,000/-

(D.D. in favour of Dean M.G.M. Medical College, Indore payable at Indore )

  - Pre Bid Meeting					: 	29.04.2013

  - Last date and time for submission of bids		: 	As per online key dates

  - Date and time of opening of Bid (Envelope “D” & “A”) : As per online key dates

  - Place of opening of bids	 		:    at http://dme.mpeprocurement.gov.in

  - Address of communication:			Chairman, Purchase Comittee & Dean,

M.G.M. Medical College,

A.B. Road, Indore (M.P.) – 452 001

- Bidders may offer their bid online for any one or more Equipments For Modular Operation Theatre. Bids will be evaluated for each equipment separately.

- Bid and bid security as specified in the bid document must be delivered as per prescribed format for online and physical submission specified in notice and tender document.

- Bids will be opened online in the presence of Bidder’s representatives who choose to attend on the specified date and time fixed for opening the bid. Envelop D containing envelop A & B  documents will be opened online as well as physically, out of which envelop A will be immediately opened. Documents together with contents of envelop A will be subject to scrutiny, those bidders whose documents and contents are as per tender conditions will only be deemed qualified for opening of technical bid. The date and time of opening of technical bid (envelop B) will be made available on website within a week of opening of bid and will be communicated  online only at http://dme.mpeprocurement.gov.in

- Price bid should be submitted online only. In case price bid is submitted manually bid will be out through rejected. Cost of tender form Rs.5,000/- (in the form of DD) should be submitted in separate Envelope marked “Tender Fee”. EMD & Technical Bid will be submitted separately in separate sealed Envelopes. Technical Bid will be evaluated as per specification and National Competitive Bidding terms and conditions by the Technical Committee. Those bidders who qualify for the technical bid will be invited for demonstration whereable applicable of equipment on the day, date & place specified by R.C.A. The bidders will have to demonstrate the equipments on the date, day and place specified. Thereafter, Price Bid (Envelope C) will be opened online only for those bidders whose bid will be found technically responsive after demonstration. Opening of date of price bid will be communicated online at  http://dme.mpeprocurement.gov.in hence  bidder are advised to visit web portal on day to day basis.

- In the event of the date specified for bid receipt/opening being declared as closed / holiday, due date for receipt/opening of bid will be post poned online.

- Supporting documents along with original EMD and cost of tender form has to be submitted by tenderer at Chairman, Purchase Committee & Dean office, M.G.M. Medical College, Indore (M.P.) on or before the time & date of submission as mentioned above.

- Time of Completion of work  within 8 weeks from the date of allotment letter.

CHAIRMAN, PURCHASE COMMITTEE &

DEAN, M.G.M MEDICAL COLLEGE,

A.B. ROAD INDORE (M.P.)

SECTION II :

INSTRUCTIONS TO BIDDER (ITB)

SECTION II: INSTRUCTIONS TO BIDDER

TABLE OF CLAUSES

| Clause No. | Topic Number | Page No. | Clause No. | Topic Number | Page No. |
| --- | --- | --- | --- | --- | --- |
|  | A. Introduction |  |  | D. Submission of Bids |  |
| 1. | Scope of work | 13 | 17. | Bid stages | 20 |
| 2. | Eligible Bidders | 13 | 18. | Deadline for Submission of Bids | 23 |
| 3. | Eligible Goods and Services | 13 | 19. | Late Bids | 23 |
| 4. | Cost of Bidding | 14 | 20. | Modification and withdrawal of bids | 24 |
|  | B. Bidding Documents |  |  | E. Bid opening and evaluation of Bids |  |
| 5. | Contents of Bidding Documents | 14 | 21. | Opening of Bids by the Rate Contracting Authority | 24 |
| 6. | Clarification of Bidding Documents | 14 | 22. | Clarification of Bids | 25 |
| 7. | Amendment of Bidding Documents | 15 | 23. | Preliminary Examination | 25 |
|  | C. Preparation of Bids |  | 24. | Evaluation and Comparison of Bids | 26 |
| 8. | Language of Bid | 15 | 25. | Contacting the Rate Contracting Authority | 26 |
| 9. | Documents Comprising the Bid | 15 |  | F. Award of Contract |  |
| 10. | Bid Form | 16 | 26. | Post qualification | 26 |
| 11. | Bid Prices | 16 | 27. | Award Criteria | 27 |
| 12. | Bid Currencies | 17 | 28. | Purchaser’s Right to vary Quantities | 27 |
| 13. | Documents establishing Bidder’s Eligibility and Qualifications | 17 | 29. | Rate Contracting Authority’s Right to Accept any Bid and to Reject any or all bids | 27 |
| 14. | Documents establishing Good eligibility and conformity to Bidding Documents | 19 | 30. | Notification of Rate Contract | 27 |
| 15. | Bid Security | 19 | 31. | Signing of Rate Contract | 27 |
| 16. | Period of Validity of Bids | 19 | 32. | Performance Security & Inspection Charges | 28 |
|  |  |  | 33. | Placement of Supply Order | 28 |
|  |  |  | 34. | Corrupt or Fraudulent Practices | 28 |

- Introduction

- Scope of Work

  - Govt. of Madhya Pradesh, Medical Education Department, requires Equipments For Establishment of Modular Operation Theatre for various Govt. / Autonomous Medical Colleges, Dental College & Associated Hospitals of Madhya Pradesh. Bid is issued for procurement of Equipments For Establishment of Modular Operation Theatre at competitive rates. After finalization of the bid, the contract will be awarded to successful bidders for supply of the items, during one year or till further order on rate contract basis on approved rates. Establishment of Modular Operation Theatre to be supplied & installed at various designated places like Indore, Bhopal, Jabalpur, Gwalior, Rewa & Sagar.

- Eligible Bidders

2.1 	This invitation for Bids is open to all eligible bidders (Please refer guidelines to bidders for online eligibility mentioned above.)

2.2 	Bidders should not be associated, or have been associated in the past, directly or indirectly with a firm or any of its affiliates which have been engaged by the Rate Contracting Authority to provide consulting services for the preparation of the design, specifications and other documents to be used for the procurement of the goods to the Rate Contracting Authority under this Invitation of Bids.

2.3 	Government owned enterprises in the Rate Contracting Authority’s country may participate only if they are legally and financially autonomous, if they operate under commercial law and if they are not a dependent agency of the Rate Contracting Authority.

2.4	The tenderers shall clarify/state whether he/they are manufacturer, accredited agent or sole representative indicating principals name & address. The offers of firms who are not manufacturer or direct authorized agent will be summarily rejected. Sub-distributors will not be accepted.

- Eligible Goods and Services

3.1 	All goods and ancillary services to be supplied under the Contract shall specify their country of origin.

3.2 	For purposes of this clause, “origin” means the place where the goods are mined, grown, or produced or from which the ancillary services are supplied. Goods are produced when, through manufacturing, processing or substantial and major assembling of components, a commercially recognized product results that is substantially different in basic characteristics or in purpose or utility from its components.

3.3	The origin of goods and services is distinct from the nationality of the Bidder.

- Cost of Bidding

4.1 	The Bidder shall bear all costs associated with the preparation and submission of its bid, and Chairman, Purchase Committee & Dean M.G.M. Medical College, Indore hereinafter referred to as “The Rate Contracting Authority” will in no case be responsible or liable for these costs, regardless of the conduct or outcome of the bidding process.

- The Bidding Documents

- Contents of Bidding Documents

5.1 	The goods required, bidding procedures and contract terms are prescribed in the Bidding Documents. In addition to the Invitation for Bids, the Bidding Documents include:

- Instructions to Bidders (ITB);

- General Conditions of Contract (GCC);

- Special Conditions of Contact (SCC);

- Bid Form;

- Annexure-I (Sales Tax Clearance Certificate);

- Annexure-II (Manufacturer Authorization Form);

- Annexure-III (Declaration / Undertaking Form);

- Annexure-IV (Proforma for Performance Statement);

- Annexure-V (Annual Turnover Statement);

- Annexure-VI (Specifications of required equipments for MODULAR OPERATION THEATRE);

- Annexure-VII (Performance Security Form);

- Annexure-VIII (Contract Agreement Form);

- Annexure-IX (Details of Manufacturing Unit);

- Annexure-X (Price Schedule); (should be quoted online only)

- Annexure-XI (Price Schedule for AMC (without spare parts) / CMC (include free labour, repair, other services & spare parts); and

- Annexure-XII (Detail of Service Centre in M.P.);

- Annexure-XIII (Check List)

5.2 	The Bidder is expected to examine all instructions, forms, terms, specifications and annexure in the Bidding Documents. Failure to furnish all information required by the Bidding Documents or submission of a bid not substantially responsive to the Bidding Documents in every respect will be at the Bidder’s risk and may result in rejection of its bid.

5.3	The biding document is not transferable.

- Clarification of Bidding Documents

6.1 	If wishes, a prospective Bidder requiring any clarification of the Bidding Documents shall contact the Rate Contracting Authority in writing at the Rate Contracting Authority’s mailing address indicated in the invitation for Bids. The Rate Contracting Authority will respond in writing to any request for clarification of the Bidding Documents, which it receives not later than 15 days prior to the deadline fixed for submission of Bids and prescribed by the Rate Contracting Authority. Any correspondence with Rate Contracting Authority seeking any clarification regarding any matter contained herein shall not compel the Rate Contracting Authority to suspend the implementation of provision given hereunder or shall not mean a promise to change any provision in this tender document.

- Amendment of Bidding Documents

7.1 	At any time prior to the deadline fixed for submission of bids, the Rate Contracting Authority may, for any justifiable reason, whether at its own initiative or in response to a clarification requested by a prospective bidder, modify the bidding documents by amendment.

7.2 	Any addendum issued shall be part of Bidding Documents and all the prospective bidders will be notified of the amendment by post or publication, and will be binding on them. The same shall be uploaded on the designated website.

7.3 	In order to allow reasonable time to prospective bidders in which to take the amendment into account in preparing their bids or for any other reason, the Rate Contracting Authority at its discretion, may extend the deadline for the submission of bids.

- Preparation of Bids

- Language of Bid

8.1 	The Bid prepared by the Bidder, as well as all correspondence and documents, printed literature and leaflets relating to the bid exchanged by the Bidder and the Rate Contracting Authority shall be written in English / Hindi language.

- Documents comprising the Bid

9.1 	The bid prepared by the Bidder shall comprise the following components:

- A Bid Form and Price Schedule completed in accordance with ITB Clause 10, 11 and 12; Price bid form should be filled online only. Technical Bid Envelope, EMD and cost of tender form instruments should be uploaded online as well as submitting hardcopies thereof.

- Documentary evidence established in accordance with ITB Clause 13 that the Bidder is eligible to bid and is qualified to perform the contract if its bid is accepted;

- Documentary evidence established in accordance with ITB Clause 14 that the goods and services to be supplied by the Bidder are eligible goods and services and conform to the Bidding Documents; and

- Bid Security furnished in accordance with ITB Clause 15.

The Bidders shall submit their Earnest Money Deposit as usual in a physically sealed Earnest Money Deposit Envelope and the same should reach the concerned office as stated in the Notice Inviting Tender by post.

The Bidders have to upload scanned copy of Earnest Money Deposit instrument along with the reference details.

- Bid Form

10.1 	The Bidder shall complete the Bid Form and shall also submit a hard copy thereof. The signing of Bid Form shall commit the Bidder to supply the ordered goods to the purchaser within 30 days of placing such order.

- Bid Prices

11.1 	Bid has been called for the equipments/machines given in the specification in Technical Annexure VI. The bidder should quote the price online only for the equipments/machine offered for. The specifications of the equipments/machines should be brand new unit as per details given in Annexure-VI. Any variation found will result in the rejection of the tender. In any case if price bid submitted manually or in CD than tender will be rejected.

11.2	Prices (inclusive of Excise Duty / Custom Duty, transportation, packing, insurance, installation, loading-unloading, warranty, service charge, inspection, and any incidental charges, but exclusive of CST/VAT) should be quoted for each of the required equipments etc., separately on door delivery basis according to the unit ordered. Tender for the supply of equipments etc. with cross conditions like “AT CURRENT MARKET RATES” shall not be accepted. Handling, clearing, transport charges etc. will not be paid. The delivery should be made as stipulated in the  supply order placed with successful bidders. Conditional tenders will not be accepted.(Please note that Price Bid/ Financial Proposal is mandate to be online only if submitted in hard copy or in CD bid will be rejected. )

  - Each bid must contain the unit price of each equipment in digits as well as alphabets. Any discrepancy between the figures and words, the amount written in words will prevail. The tenders should be digitally signed online. The tenderers should take care that the rates and amounts are written in such a way that interpolation is not possible, no blanks should be left which would otherwise, make the tender redundant.

  - The price quoted by the bidders shall not, in any case exceed the controlled price, if any, fixed by the Central/State Government and the Maximum Retail Price (MRP). Rate Contracting Authority at its discretion, will exercise, the right to revise the price at any stage, on lower side so as to confirm to the controlled price or MRP as the case may be. This discretion will be exercised without prejudice to any other action that may be taken against the bidder.

  - To ensure sustained supply without any interruption the Rate Contracting Authority, reserves the right to split orders for supplying the requirements among more than one bidder, provided the prices and other conditions of supply are equal.

  - The prices quoted and accepted will be binding on the bidder for the stipulated period (as per para I of introduction) and any increase in the price will not be entertained till the completion of this tender period or till further orders. Cross Conditions such as “SUBJECT TO AVAILABILITY” “SUPPLIES WILL BE MADE AS AND WHEN SUPPLIES ARE RECEIVED” etc., will not be considered under any circumstances and the tenders of those who have given such conditions shall be treated as incomplete and Tender will be summarily rejected.

  - The price quoted should be inclusive of service tax (as per Govt. rules) on inspection and testing charges, which will be realized in advance from the bidders.

- Bid Currencies

12.1 	Prices shall be quoted in Indian Rupees.

- Documents establishing Bidder’s eligibility and qualifications

13.1 	Pursuant to ITB Clause 9, the Bidder shall furnish, as part of its bid, documents establishing the bidder’s eligibility to bid and its qualifications to perform the Contract if its bid is accepted.

13.2 	The documentary evidence of the Bidder’s eligibility to bid shall establish to the Rate Contracting Authority’s satisfaction.

13.3 	The documentary evidence of the Bidder’s qualifications to perform the Contract if its bid is accepted, shall establish to the Rate Contracting Authority’s satisfactions:

- Bids may be submitted by the primary manufacturer or their authorized distributor or importer for and on behalf of the primary manufacturer provided the bid is accompanied by a duly notarized letter of authority from the primary manufacturer. In case of authorized distributor the bidder should have minimum three years association with manufacturer. (as per authorization form given in Annexure II).

- Documentary evidence for the Registration of the company with details of the Name, Address, Telephone Number, Fax Number, e-mail address of the firm and of the Managing Director / Partners / Proprietor and Name, Address, Telephone number, fax, e-mail of primary manufacturer.

- The bidder shall submit printed original catalogues of primary manufacturer and any other technical documents like data sheet or operational manual of equipment with highlighting the features in portal along with the other documents. In catalogue, the quoted product no. and name should be highlighted and item code should also be written with catalogue, against which that product is quoted. These documents are also to be submitted in physical form before due date along with Bid security. Specification of equipments supplied should match the specification in catalogue.

- The instruments such as power of attorney, resolution of board etc., authorizing an officer/person of the bidder should be submitted with the tender and such Authorized officer/person of the bidder should sign the tender documents.

- Authorization letter nominating a responsible person of the bidder to transact the business with the Rate Contracting Authority.

- The Bidder/manufacturer should have atleast three years manufacturing / distributorship experience. The Bidder should submit a list of user of quoted equipments manufactured by the Principal Manufacturer for last three years. These list should also contain the supplies related to the Govt. hospital / Medical Colleges / Public Sector undertaking / Undertaking hospital and other institutions of repute. Bidder should submit details of installation in Annexure IV.

- The bidder should have at least one service centre in Madhya Pradesh, with a team of trained service engineer/technical staff the details in this regard as per Annexure-XII shall be submitted. In case at the time of tender service centre is not available in M.P., then he shall submit undertaking to establish the service centre before the award of contract. It shall be the duty of the seller to collect the equipment for repair / service & to replace it after such repair / servicing free of charge during warranty / guarantee period and will cover all such costs in the AMC / CMC.

- The bidder shall submit the specification’s compliance / deviation report duly filled and signed which clearly bring out the deviation from the specification if any given in Annexure-VI.

- Sales Tax/VAT/CST Clearance certificate, as on 31.03.2010 / 31.03.2011 / 31.03.2012 (as per form attached in Annexure-I).

- Details of Manufacturing Unit I Annexure – IX. The details containing the name and address of the premises where the items quoted are actually manufactured.

- Documents, if any, to show that the manufacturing unit/importer has been recognized, by WHO, UNICEF, ISO or any other Certificate etc.

- The bidder shall furnish a notarized affidavit in the format given in Annexure-III declaring that the bidder accepts all terms and conditions of the tender.

- Annual turnover (i.e. turnover for each year separately) in the last three financial years shall not be less than Rs. Three Crore for Manufacturer and Rs. One Crore Fifty Lacs for the authorized distributor. Annual turnover statement for 3 years submitted in the format given in Annexure-V certified by the Auditor/CA.

- In case of imported equipment IEC certificate of importer / bidder shall be submitted.

- The bidder should also submit national and international quality certificates like ISI/CE/C” mark/IEC standard or equivalent certificate of quoted product, if available.

- Concern / Company have not been debarred / blacklisted either by Rate Contracting Authority or by any State Government or Central Government Organization for the quoted product or as a whole. Affidavit to this effect shall be submitted by the concern / company.

- Leaflets, literatures, should invariably be attached for ready reference clearly marking the item code no.

- All documents should be self attested and stamped.

  - The work done by the applicant company in India whether private or Government sector will only be considered for evaluation process.

  - The company engaged in Modular Operation Theatre manufacturing should have minimum five years of market standing, in India. Relevant evidentiary documents should be submitted in support thereof.

  - The applicant company should have exposure to complete project of manufacturing / installation / fixing of Modular Operation Theatre work at at-least 5 Government / Non Government Medical College / Medical University / Hospitals over the past 5 years in India.

  - No Bidder / Tenderer should submit two or more authorization letters from the manufacturer for the same projects / tender.

  - The applicant company should be ISO & CE certified.

- Documents establishing Goods Eligibility and Conformity to Bidding Documents

14.1 	Pursuant to Clause 9, the Bidder shall furnish, as part of its bid, documents establishing the eligibility and conformity to the Bidding Documents of all goods and services, which the Bidder proposes to supply under the Contract.

- Bid Security

15.1 	Pursuant to ITB Clause 9, the Bidder shall furnish, as part of its bid, the bid security which shall be Rs. 2.00 Lakhs (Rs. Two Lakhs). No concession / exemption shall be allowed. F.D.R. should be attached. Scan Copy of same instrument should be uploaded online too.

15.2 	The bid security is required to protect the Rate Contracting Authority against risk of Bidder’s business conduct which would warrant the security’s forfeiture, pursuant to ITB clause 15.7.

15.3 	The bid security shall be in Indian Rupees and shall be in form of Fixed Deposit for 18 months in favour of Dean, M.G.M. Medical College, Indore.

15.4 	Any bid not secured in accordance with ITB Clause 15.1 and 15.3 above will be rejected by the Rate Contracting Authority as non-responsive, pursuant to ITB Clause 23.

15.5 	Unsuccessful Bidder’s bid security will be discharged / returned as promptly as possible upon the successful Bidders signing the Contract, pursuant to ITB Clause 31 or after the expiration of the period of bid validity prescribed by the Rate Contracting Authority pursuant to ITB clause 16. No interest is payable on bid security.

15.6 	The successful Bidder’s bid security will be discharged upon the Bidders signing the Contract, pursuant to ITB Clause 30, and furnishing the performance security, pursuant to ITB Clause 31. No interest is payable on bid security.

15.7 	The bid security may be forfeited:

a. 	If a bidder

(i) withdraws its bid during the period of bid validity specified by the Bidder on the Bid Form;

b.	In case of a successful Bidder, if the Bidder fails:

(i) to sign the Contract in accordance with ITB Clause 30; or

(ii) to furnish performance security and Inspection Charges in accordance with ITB Clause 31.

- Period of Validity of Bids

16.1 	Quoted Prices of Bids shall be valid for 180 (One hundred eighty) days after the date of bid opening prescribed by the Rate Contracting Authority pursuant to ITB clause 21. A bid valid for a shorter period shall be rejected by the Rate Contracting Authority as non-responsive. This price  on acceptance shall remain fixed till contract period or till further order.

16.2 	In exceptional circumstances, the Rate Contracting Authority may solicit the Bidder’s consent to an extension of the period of validity. The request and the responses thereto shall be made in writing. A Bidder may refuse the request without forfeiting its bid security. A bidder granting the request will not be required nor permitted to modify its bid.

16.3 	No bid may be modified subsequent to the deadline for submission of Bids.

16.4 	No Bid may be withdrawn in the interval between the deadline for the submission of Bids and expiration of the period of bid validity specified by the Bidder on the Bid Form. Withdrawal of a bid during this interval may result in the Bidders forfeiture of its bid security pursuant to ITB Sub-clause 15.7.

- Submission of Bids

- Bid Stages

17.1	Bid should be submitted in following system and should furnish the following documents failing which their bid shall not be accepted:-

Envelope – A:  Tender fee and E.M.D.

(Tender fee in the form of D.D. only)

- The Bidders shall submit their Tender fee Rs.5,000/- in form of a Demand Draft in the name of “Dean, M.G.M. Medical College, Indore” payble at Indore in a separate physically sealed Envelope clearly marked “Tender Fee”and the same should reach the concerned office as stated in the Notice Inviting Tender by post. The Bidders have to upload scanned copy of the Tender fee in form of Demand Draft instrument along with the reference details..

(E.M.D. in the form of F.D.R. only)

(2)       The Bidders shall submit their Earnest Money Deposit as usual in a separate

physically sealed Envelope and the same should reach the concerned office as stated in the Notice Inviting Tender by post. The Bidders have to upload scanned copy of Earnest Money Deposit instrument along with the reference details..

- Fixed Deposit for 18 Months as bid security which shall be Rs. 2.00 Lakh (Rs. Two Lakh only) tendered for as E.M.D. issued by Nationalized bank in the name of the Dean, MGM Medical College, Indore, payable at Indore (refundable). (Any other example D.D. / Cheque will not be accepted).

- Without submission of E.M.D. the tender will be summarily rejected as per rules.

- In no case the tender cost fee should be mixed with E.M.D. amount. Fee cost is not refundable.

Envelop – B: (Technical Bid)

- Technical bid should be submitted online as well as with two additional self certified copies (total 3 copies) for the quoted equipments etc. should be signed and stamped on each page. (ANNEXURE-VI). The bidder shall submit the specification’s compliance / deviation report duly filled and signed which clearly bring out the deviation from the specification if any given in Annexure-VI. Format of technical bid is available in tender document and bidder should provide technical bid in MS office version 97 in C.D. and has to be submitted along with the hard copy by the bidder in Envelope B.

- List of name and address where supply of the quoted equipments has been made.

- Literature of original catalogue of the product attached for reference.

- Guarantee / warrantee.

- Sales Tax/VAT/CST Clearance certificate, as on 31.03.2010, 31.03.11, 31.03.12 (as per form attached in Annexure-I).

- Annexure-II (Manufacturer Authorization Form)

- Annexure-III (Undertaking Form / Declaration Form)

- Annexure-IV (Proforma for Performance Statement). A list of user of quoted equipments by the Principal Manufacturer for last three years. These list should also contain the supplies related to the Govt. hospital / Medical Colleges / Public Sector undertaking / Undertaking hospital and other institutions of repute. Bidder should submit details of installation in Annexure IV.

- Annexure-V (Annual Turnover Statement)

- Annexure – IX (Details of Manufacturing Unit)

- Annexure – XII (Details of Service Centre in M.P.)

- Registration Certificate of the company with details of the Name, Address, Telephone Number, Fax Number, e-mail address of the firm and of the Managing Director / Partners / Proprietor.

- Authorization letter from manufacturer authorizing a person to transact a business with R.C.A.

- The instruments such as power of attorney, resolution of board etc., authorizing an officer/person of the bidder should be submitted with the tender and such Authorized officer/person of the bidder should sign the tender documents.

- Market Standing Certificate issued by the Licensing Authority as a Manufacturer / distributor for each equipment quoted for the last 3 years. In case of direct importer, evidence for importing the said items for the last three years.

- The bidder should also submit national & international quality certificates like ISI/CE/C ISO-9002, IP/BP etc” mark / IEC standard or equivalent certificate of quoted product, if available.

- Concern / Company have not been debarred / blacklisted either by Rate Contracting Authority or by any State Government or Central Government Organization. Affidavit to this effect shall be submitted by the concern / company.

- Original price Bid and other Form duly digitally signed online by authorized signatory and physical document as per Section V, duly sealed and signed by the bidder on each page for acceptance of Terms and Conditions.

- Bidders should have the registration under Commercial Tax Authority, Registration should be attached.

- Affidavit that the firm has no vigilance case / CBI case pending against him / supplier.

- Affidavit that the firm has not supplied the same item at the lower rate than quoted in the tender to any Govt. / Semi Govt. or any other organization.

- Certificate for being in business or more than 3 years.

- Certificate for sole ownership / partnership and establishment relationship.

- Statement of good financial standing from bankers.

- The printed original catalogues of primary manufacturer and any other technical documents like data sheet or operational manual of equipment with highlighting the features in portal along with the other documents. In catalogue, the quoted product no. and name should be highlighted, against which that product is quoted. These documents are also to be submitted in physical form before due date along with Bid security.

- In case of imported equipment IEC certificate of importer / bidder shall be submitted.

- A separate price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee / warrantee period.

- Recurring expenditure on equipments.

Both envelops sealed in main envelop i.e. marked ENVELOP–D   “TENDER FOR ESTABLISHMENT OF MODULAR OPERATION THEATRE FOR P.C.” All the envelopes A, B, D must be Wax sealed using sealing Wax and official seal, sealed cello taped, moisture free and strong. All the enclosures and photocopies should be self certified and stamped.

    - Reference No. of the tender ________________

    - Tender regarding _________________________

    - Due date of submission of tender form ________

    - Due date for opening of the tender ___________

    - Name of the firm _________________________

NOTE : TENDER SUBMITTED WITHOUT FOLLOWING THE ABOVE PROCEDURES WILL BE SUMMARILY REJECTED.

Envelop – C: (Financial Bid)

Financial Bid  / Price Proposal should be submitted online only, in case  if the price bid is sent physically or in  CD as a soft copy the bid will be out through rejected.

Envelop – D: (Main Tender Envelop)

  - Envelop A & B.

  - PRICE BID –

The Bidder should furnish the following:-

- The rate quoted online only as per unit (landed price) in Annexure-X (Suggested sample proforma of price schedule in this tender document is for inspection and not for physical submission or in CD) shall be inclusive of Excise Duty / Custom Duty, freight, packing, insurance, inspection & testing charges etc. exclusive of VAT and should be F.O.R. point of installation on turn key basis.

- The rate quoted in column 8 of Annexure-X should be filled online(Suggested sample proforma of price schedule in this tender document is for inspection and not for physical submission or in CD) should be for a unit and given specification. The Bidder is not permitted to change / alter specification or unit size given in the Annexure-X.

- Bidder has to quote rates online only strictly for the items which are mentioned in the tender.

- Rates quoted for items other than mentioned in the tender form then that particular item will not be entertained.

- The rates of each item should be quoted in figures as well as in words also otherwise the tender is liable to be rejected.

- The bidder shall also quote charges for Annual Maintenance Contract (without spare parts) / Comprehensive Maintenance Contract (include free labour, repair, other services & spare parts) for the next five years after the expiry of five years warranty/gurantee period in Annexure-XI. AMC should be quoted for equipments costing upto Rs.5.00 Lacs and CMC should be quoted for equipments costing more than Rs.5.00 Lacs.

- The bidder should quote equipments which will have guarantee / warranty of atleast 5 years, equipments which have less than 5 years warranty will not be entertained and so should not be quoted.

- Bidder should show recurring expenditure of each equipment separately.

- The price bid should be submitted online as prescribed in online form only. In case price bid / financial bid is submitted manually or in CD than bid will be rejected.

- Deadline for Submission of Bids

  - Bids will not be accepted after the time and date specified in the invitation for Bids (Section I).

  - The Rate Contracting Authority may, as its discretion, extend the deadline for submission of bids by amending the Bid Documents in accordance with ITB Clause 7, in which case all right and obligations of the Rate Contracting Authority and Bidders previously subject to the deadline shall thereafter be subject to the deadline as extended.

  - The Rate Contracting Authority will not be responsible for any delay or non-receipt of tender documents.

- Late Bids

  - No Bid can be submitted after the last date and time of submission of bid.

- Modification and withdrawal of Bids

  - The Bidder may modify or withdraw their bid before Last Date and time of submitting bid hash online, before completing Submit Bid Hash Online stage. Once the said stage shows to be completed, no modifications can be made by bidder.

  - No bid may be modified subsequent to the deadline fixed for submission of bids hash online.

  - No bid may be withdrawn in the interval between the deadline for submission of bids and the expiration of the period of bid validity specified by the Bidder on the Bid Form or any extension thereof. Withdrawal of a bid during this interval may result in the Bidder’s forfeiture of its bid security, pursuant to ITB Clause 15.7.

- Bid Opening and Evaluation of Bids

- Opening of Bids by the Rate Contracting Authority

Opening of Bid process:

  - Bid will be opened online as per key schedule on the day, date & place specified by the Rate Contracting Authority.

  - All bidders are entitled to be present at the date, time & place for opening of Bids.

  - Only one representative of each Bidder is entitled to remain present at the time of bid opening. Bidder’s representative who is present shall sign a register evidencing his/her attendance. In the event of the specified date of Bid opening being declared a holiday for the Rate Contracting Authority, the Bids shall be opened at the appointed time and location on the next working day. Such representative must be present with original authority letter issued by the bidder with ID proof in support thereof.

  - Opening of bid will be sequential process.

  - Bids will be opened in the presence of Bidder’s representatives who choose to attend on the specified date and time fixed for opening the bid. Envelop D containing envelop A & B documents will be opened, out of which envelop A will be immediately opened online and physically. Documents together with contents of envelop A will be subject to scrutiny, those bidders whose documents and contents are as per tender conditions will only be deemed qualified for opening of technical bid. The date and time of opening of technical bid (envelop B) will be made available on website within a week of opening of bid and will be communicated simultaneously to those who qualify for opening of bid. Technical and Price bid will be submitted online and envelope B will be submitted in hard copy too. Technical Bid will be evaluated as per specification and NCB terms and conditions by the Technical Committee. Those bidders who qualify for the technical bid will be invited for demonstration of equipment on the day, date & place specified by R.C.A. The bidders will have to demonstrate the equipments on the date, day and place specified, failing which their bid will be rejected. Thereafter, Price Bid (envelop C) will be opened online only for those bidders whose bid will be found technically responsive after demonstration. Opening of date of price bid will be communicated online only at http://dme.mpeprocurement.gov.in those who qualify and will be displayed on the specified website.

  - The Bidders’ names, presence or absence of the requisite bid security will be announced at the opening of Technical Bid.

  - Bidders who were found eligible on satisfying the criteria for technical evaluation and inspection by the technical committee can only be invited to be present at the date and time for opening of Price Bid of the tender.

- Clarification of Bids

  - During evaluation of bids, the Rate Contracting Authority may, at its discretion, ask the Bidder for clarification of its Bid. Any clarification submitted by a bidder in respect to its bid and that is not in response to a request by the Rate Contracting Authority shall not be considered. The request for clarification and the response shall be in writing and no change in prices or substance of the bid shall be sought, offered or permitted except to confirm the correction of arithmetic errors discovered by the Rate Contracting Authority in the evaluation of the bids.

- Preliminary Examination

  - The Rate Contracting Authority will examine the bids to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether all documents are there, whether the documents have been properly signed, and whether the bids are generally in order.

  - The Rate Contracting Authority may waive any minor informality or non-conformity or irregularity or omissions in a bid which does not constitute a material deviation, provided such a waiver does not prejudice or affect the relative ranking of any Bidder.

  - Prior to the detailed evaluation, pursuant to ITB Clause 24, the Rate Contracting Authority will determine the substantial responsiveness of each bid to the bidding documents. For purposes of these Clauses, a substantially responsive bid is one which conforms to all the documents, terms, conditions and specifications of the bidding documents without material deviations. The Rate Contracting Authorities determination of a bid’s responsiveness is to be based on the contents of the bid itself without recourse to extrinsic evidence.

  - If a bid is substantially responsive, the Rate Contracting Authority may request that the bidder submit the necessary information or documentation, within a reasonable period of time to rectify nonmaterial nonconformities or omissions in the bid related to documentation requirements. Such omissions shall not be related to any aspect of the price of the bid. Failure of the bidder to comply with the request within the stipulated time may result in the rejection of its bid.

  - If a bid determined as not substantially responsive, it will be rejected by the Rate Contracting Authority and may not subsequently be made responsive by the bidder by correction of the non-conformity.

- Evaluation and Comparison of Bids

  - The Rate Contracting Authority will evaluate and compare the bids previously determined to be substantially responsive, pursuant to Clause 23. Bids will be evaluated with reference to various criteria as specified in bid document and one of such criteria is that the rate per unit of (landed price) i.e. rate per item  for determining the L1 rate (Lowest rate).

  - Purchase will also be made from SC / ST firms as per Madhya Pradeh State Government rules.

- Contacting the Rate Contracting Authority

  - Subject to ITB Clause 22, no Bidder shall contact the Rate Contracting Authority on any matter relating to its bid, from the time of the bid opening to the time Rate Contract is awarded.

  - Any effort by a Bidder to influence the Rate Contracting Authority in its decisions on bid evaluation, bid comparison or contract award may result in rejection of the Bidder’s bid. If the bidder wishes to bring additional information to the notice of the Rate Contracting Authority, it should do so in writing.

- Award of Contract

- Post Qualification

  - Based on the qualification criteria listed in ITB Clause 13, the Rate Contracting Authority will determine to its satisfaction whether the Bidder selected as having submitted the lowest evaluated responsive bid is qualified to satisfactorily perform the Contract.

  - The determination will take into account the Bidder’s financial, technical, and production capabilities. It will be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted by the Bidder. Pursuant to ITB Clause 13, as well as such other information as the Rate Contracting Authority deems necessary and appropriate.

  - The Rate Contracting Authority shall ask for demonstration whereable applicable of the quoted Modular Operation Theatre equipment. The cost of demonstration shall be born by the bidder. Day, date & place of demonstration shall be decided by Rate Contracting Authority.

  - An affirmative determination will be prerequisite for award of the Rate Contract to the Bidder. A negative determination will result in rejection of the Bidder’s bid in which event the Rate Contracting Authority will proceed to the next bid to make a similar determination of that Bidder’s capabilities to perform the contract satisfactorily.

  - The tenderers shall demonstrate whereable applicable the quoted model of the equipments during the technical evaluation on the day, date & place specified.

- Award Criteria

  - Subject to ITB Clause 29 – The Rate Contracting Authority will award rate contract to the successful bidders on lowest evaluated prices or the price approved by the Purchase Committee.

- Purchaser’s Right to vary Quantities

  - The details of the required equipments etc. are shown in Annexure-VI. The quantity mentioned is only the tentative requirement and may increase or decrease as per the decision of the Purchaser. The rates quoted should not vary with the quantum of the order or the destination.

- Rate Contracting Authority’s Right to Accept any Bid and to Reject any or all bids

  - The Rate Contracting Authority reserves the right to accept or reject the tender for the supply of all items of equipments or for any one or more of the items of equipments tendered for in a tender without assigning any reason, without thereby incurring any liability to the affected Bidder or Bidders or any obligation to inform the affected Bidder or Bidders of the grounds for the Rate Contracting Authority’s action.

- Notification of Rate Contract

  - Prior to the expiration of the period of bid validity, the Rate Contracting Authority will notify the successful Bidder in writing by registered letter or fax or e-mail, that its bid has been accepted.

  - The notification of Rate Contract will constitute the formation of the Contract.

  - Upon the successful Bidder’s signed Rate Contract pursuant to ITB Clause 31, the Rate Contracting Authority will promptly notify each unsuccessful Bidder and will discharges its bid security, pursuant to ITB Clause 15.

  - If, after notification of rate contract, a Bidder wishes to ascertain the grounds on which its bid was not selected, it should address it’s request to the Rate Contracting Authority. The Rate Contracting Authority will promptly respond in writing to the unsuccessful Bidder.

- Signing of Rate Contract (Agreement)

  - At the same time the Rate Contracting Authority will inform to the successful Bidder that its bid has been accepted the Rate Contracting Authority will send the Bidder the Rate Contract Form provided in the bidding document incorporating all agreements between the parties.

  - Within 10 days of receipt of the Notification of Rate Contract, the successful Bidder shall sign and date the Contract on a non-judicial stamp paper of value of Rs.100/- (stamp duty to be paid by the Bidder) and return it to the Rate Contracting Authority.

  - The validity of Rate Contract will be one year and may be extended for further period as agreed mutually unless revoked.

- Performance Security & Inspection Charges

  - Within 15 days of the receipt of firm order from the Rate Contracting Authority or the date specified by the purchaser, the successful Bidder shall furnish the performance security and inspection charges in accordance with the Clause 7 & 8 of General Conditions of Contract.

  - Failure of successful bidder to comply with the requirement of ITB Clause 31 or ITB Clause 32.1 shall constitute sufficient grounds for the annulment of the award and forfeiture of the bid security/previously deposited performance security and disqualify the firm to participate in the tender for the next five years.

- Placement of Supply Order

  - After finalization of the contract, the successful bidders may be asked to submit the delivery schedule as per requirement of the Purchaser. While placement of orders, the schedule given to the bidders, along with the other conditions stated at ITB 27.1 will be considered.

  - To ensure sustained supply without any interruption the Purchaser, reserves the right to split orders for supplying the requirements among more than one L-1 bidder.

- Corrupt or Fraudulent Practices

  - For the purpose of this provision, the terms set forth as follows:

- “Corrupt practice” means the offering, giving, receiving or soliciting of any thing of value to influence the action of a public official in the procurement process or in contract execution, and

- “Fraudulent practice” means a mis-presentation / hiding of facts in order to influence a procurement process or the execution of a contract to the detriment of the other bidders, and includes collusive practice among Bidders (prior to or after bid submission) designed to establish bid prices at artificial noncompetitive levels and to deprive the other bidders of the benefits of free and open competition;

- Will reject a proposal for award if it determines that the Bidder recommended for award has engaged in corrupt or fraudulent practice in competing for the contract in question.

- Will declare a firm ineligible, either indefinitely or for a stated period of time, to be allowed to participate, awarded a contract if at any time determines that the firm has engaged in corrupt or fraudulent practice in competing for, or in executing, a contract.

  - Furthermore, Bidders shall be aware of the provision stated in sub clause 21.4 of the General Conditions of contract.

SECTION III:

GENERAL CONDITIONS OF CONTRACT (GCC)

SECTION III:

GENERAL CONDITIONS OF CONTRACT (GCC)

TABLE OF CLAUSES

| Clause Number | Topic | Page Number |
| --- | --- | --- |
| 1. | Definitions | 31 |
| 2. | Applications | 31 |
| 3. | Country of Origin | 32 |
| 4. | Standards | 32 |
| 5. | Use of Contract documents and information | 33 |
| 6. | Patent Rights | 33 |
| 7. | Performance Security | 33 |
| 8. | Inspection and Tests | 34 |
| 9. | Packing | 35 |
| 10. | Delivery and Documents | 35 |
| 11. | Insurance | 35 |
| 12. | Transportation | 35 |
| 13. | Warranty | 36 |
| 14. | Payment | 38 |
| 15. | Prices | 38 |
| 16. | Change orders | 38 |
| 17. | Contract Amendments | 39 |
| 18. | Assignment | 39 |
| 19. | Delays in the Supplier’s Performance | 39 |
| 20. | Liquidated Damages | 39 |
| 21. | Termination for Default | 40 |
| 22. | Force Majeure | 41 |
| 23. | Termination for insolvency | 41 |
| 24. | Termination for Convenience | 41 |
| 25. | Resolution of Disputes | 42 |
| 26. | Limitation of Liability | 42 |
| 27. | Governing Language | 42 |
| 28. | Applicable Law | 43 |
| 29. | Notices | 43 |
| 30. | Taxes and Duties | 43 |
| 31. | Fall Clause | 43 |
| 32. | Jurisdiction | 44 |

General Conditions of Contract

  - Definitions

  - In this Contract, the following terms shall be interpreted as indicated:

- “Rate Contract” means the agreement entered into between the Rate Contracting Authority and the Supplier, as recorded in the Contract Forms signed by the parties, including all the attachments and appendices thereto and all documents incorporated by reference therein for supply of material in agreed time period.

- “Price” means the price payable to the Supplier for the full and proper performance of its contractual obligations.

- “Goods” means all the equipments etc., which the supplier is required to supply to the purchaser under the Contract.

- “Services” means services ancillary to the supply of the Goods, such as transportation and insurance and any other incidental services, and other obligations of the Supplier covered under the Contract.

- “GCC” means the General Conditions of Contract contained in this section.

- “SCC” means the Special Conditions of Contract.

- “The Purchaser” means the The Deans of Govt. Medical Colleges and Associated Hospitals of Indore, Bhopal, Gwalior, Jabalpur, Rewa & Sagar purchasing the goods, as named in SCC.

- “The Purchaser’s Country” is the country named in SCC.

- “The Supplier” means the individual or firm supplying the Goods and Services under this Contract.

- “Chairman, PC” means Chairman, Purchase Committee which is Dean, M.G.M. Medical College, Indore

- “Rate Contracting Authority” means the Chairman, Purchase Committee.

- “The Project Site” where applicable, means the place or places named in SCC.

- “Day” means calendar day.

  - Applications

  - These General Conditions shall apply to the extent that they are not superseded by provisions in other parts of the Contract.

  - Country of Origin

  - All goods and services supplied under the Contract shall be specified their country of origin.

  - For purpose of this Clause “origin” means the place where the Goods are mined, grown or product, or from which the Services are supplied. Goods are produced when, through manufacturing, processing, or substantial and major assembling of components, a commercially recognized new product results that is substantially different in basic characteristics or in purpose or utility from its components.

  - The origin of Goods and Services is distinct from the nationality of the Supplier.

  - Standards

  - The Goods supplied under this Contract shall confirm to the standards mentioned in the Technical Specifications and when no applicable standard is mentioned, latest standards agreeable to Rate Contracting Authority should be supplied.

  - Genuine Equipments For Establishment of Modular Operation Theatre must be supplied. Tenderers should indicate the source of supply i.e. name & address of the manufacturers from whom the items are to be imported.

  - While quoting the rates of Equipments of establishment of Modular Operation Theatre as enclosed list, the name of the manufacturer, must be mentioned otherwise the tender is liable to be rejected.

  - The rates of every item should be quoted from standard and well reputed firms / companies and they should be minimum possible.

  - For Equipments for establishment of Modular Operation Theatre means should bear quality assurance certification like ISO 9002 of CE Mark of ISI standardization.

  - Software and Hardware Upgradation – Free Digital Up-gradation of software (all update & upgrades) upto 5 years.

  - Voltage stabilizer & digital technology should be supplied with the equipments required it.

  - Technical specification of equipments / work mentioned is basic, however, equipments of higher specifications may be quoted at no extra cost.

  - No change in make/manufacturer will be allowed at the time of supply. Changes resulting out of technology upgradation of the same manufacturer can be permitted at no extra cost.

  - Circuit diagram with operator’s and service manual must be enclosed along with the equipment.

  - Names of the institution in India, where quoted equipment / work has been supplied / installed / done during last three years must be attached. Also number of units sold in India must be informed in writing.

  - The Bidders are not allowed to quote for equipments / components with less than desire specification. Deviation from specification on lower / negative side shall not be considered if at any time during evaluation / after supply of equipments / components are found below specification EMD / performance guarantee shall be forfeited and action will be taken for black listing.

  - Latest models which fulfills this tender’s specifications must be quoted.

  - An affidavit of the manufacturer duly notarized on Non Judicial stamp paper must be enclosed to guarantee supply of all spare parts for 5 years beyond guarantee / warranty period must be enclosed. The manufacturer must submit an authorized price list of genuine spare parts / standard parts and must also mention in the above affidavit that they will ensure sending any revised price list in the event of award of contract which requires AMC / CMC.

  - Tenders of  refurbished equipments / machineries will not be accepted. The bidder must give an affidavit on a duly notarized Non judicial stamp paper that the quoted  equipment / machine is not refurbished.

  - Use of Contract documents and information

  - The supplier shall not, without the Rate Contracting Authority’s prior written consent, disclose the Contract, or any provision thereof, or any specification, plan, drawing, pattern, sample or information furnished by or on behalf of the Rate Contracting Authority in connection therewith, to any person other then a person employed by the supplier in performance of the contract. Disclosure to any such employed person shall be made in confidence and shall extent only so far as may be necessary, for purposes of such performance.

  - The Supplier shall not, without the Rate Contracting Authority’s prior written consent, make use of any document or information enumerated in GCC Clause 5.1 except for the purposes of performing the Contract.

  - Any document, other than the Contract itself, enumerated in GCC Clause 5.1 shall remain the property of the Rate Contracting Authority and shall be returned (in all copies) to the Rate Contracting Authority on completion of the Supplier’s performance under the contract if so required by the Rate Contracting Authority.

  - Patent Rights

  - The Supplier shall indemnify the Rate Contracting Authority against all third-party claims of infringement of patent, trademark or industrial design rights arising from use of the Goods or any part thereof in India.

  - Performance Security

  - The supplier shall furnish performance security in the amount specified in SCC 2.1 to the purchaser as specified in GCC 1.1 (g).

  - The proceeds of the performance security shall be payable to the purchaser as compensation for any loss resulting from the Supplier’s failure to complete its obligations under the Contract.

  - The Performance Security shall be denominated in Indian Rupees and shall be in the form of F.D.R. / Bank Guarantee of Nationalized Bank located in India in the prescribed form provided in bidding document or another acceptable to the purchaser in favour of Purchaser till completion of warranty period.

  - The performance security will be discharged by the purchaser and returned to the Supplier not later than 60 days following the date of completion of the Supplier’s satisfactory performance obligations, including warranty obligations, unless specified otherwise in SCC.

  - Inspection and Tests

If purchaser wishes:

  - The purchaser or its representative shall have the right to inspect and/or test the Goods to confirm their conformity to the contract. The Special Conditions of Contract and/or the Technical Specification shall specify what inspections and tests the purchaser requires and where they are to be conducted. The purchaser shall notify the Supplier in writing of the identity of any representatives retained for these purposes.

  - The Supplier shall notify the purchaser or its representative at least 10 days prior to the date when Goods are available for inspection.

  - The Supplier will provide to the purchaser or its representative all reasonable facilities for the conduct of such inspections and tests at no additional cost to the purchaser. The Supplier may seek an independent quality test report for batch ready for shipment. The cost of such tests will be borne by the Supplier.

  - Where the Supplier contests the validity of the rejection by the purchaser or his representative, whether based on product or packing grounds, a sample drawn by the Inspection Authority will be forwarded for analysis to an independent technical inspection. The Finding, which will be promptly obtained, will be final and biding on both the parties. The cost of umpire analysis will be borne by the losing party.

  - The Purchaser’s right to inspect, test and where necessary, reject the Goods after the Goods arrival in at Site shall in no way be limited or waived by reason of the Goods having previously been inspected, tested and passed by the purchaser or its representative prior to the Goods shipment from the country of origin.

  - Nothing in Clause 8 shall in any way release the supplier from any warranty or other obligations under this Contract.

  - Packing

  - The Supplier shall provide such packing of the Goods as is required to prevent their damage or deterioration during transit to their final destination as indicated in the contract. The packing shall be sufficient to withstand, without limitation, rough handling during transit and exposure to extreme temperatures, salt and precipitation during transit and open storage. Packing case size and weights shall take into consideration, where appropriate the remoteness of the Goods final destination and the absence of heavy handling facilities at all points in transit.

  - The packing marking and documentation within and outside the packages shall comply strictly with such special requirements as shall be provided for in the contract including additional requirements, if any, specified in SCC, Technical specification and in any subsequent instruction ordered by the Purchaser.

  - Delivery and Documents

  - The supply should be completed within 30 days from the date of supply order unless otherwise specified in the supply order. Purchaser will place order by fax &/or e-mail &/or speed post.

  - It shall be the responsibility of the Supplier to make good for any shortage/damage at the time of receipt at designated place.

  - The details of shipping and/or other documents to be furnished by the Supplier are specified in SCC.

  - The delivery of Equipments for Modular Operation Theatre should be made at the point / place specified by the Purchaser in Purchase Order.

  - The successful bidders should strictly adhere to the following delivery schedule Supply, Installation & Commissioning on turn key basis should be effected within a fortnight from the date of supply and this clause should be strictly adhere to failing which necessary administrative action as deemed fit under rules will be taken against the defaulter.

  - Supply must be toto i.e. not in fraction.

  - Insurance

  - The Goods supplied under the contract shall be fully insured in Indian Rupees against the loss or damage incidental to manufacture, acquisition, transportation, storage, delivery, installation and test running in the manner specified in SCC.

  - Transportation

  - Where the Supplier is required under the Contract to transport the Goods to project site, including insurance as shall be specified in the Contract shall be arranged by the Supplier, and the related cost shall be included in the Contract Price.

  - The loss or damage of material whatsoever, whether insured or not, during transit shall be made good by bidder free of charge, failing which the losses will be deducted from their bill / performance security.

  - Wharf age, demurrages etc. on account of incorrect or delayed dispatch of material or documents shall be the responsibility of the supplier and shall be recovered from his bill / performance security.

  - Warranty

  - The Bidder shall provide on site warranty/gurantee of the equipment for the period of five years from the date of installation. Warranty will cover services, repairs, maintenance, replacement of spare parts, broken / damaged / worn out spare parts and other services free of cost during the whole warranty period of five years. The warranty shall also include “on call service” which should not exceed 48 hours from the date of lodging of complaint. The purchaser shall have the right to get the work done at the cost of bidder’s responsibility, if machine is not repaired within 48 hours.

  - The Purchaser shall promptly notify the Supplier in writing of any claims arising under the warranty.

  - Upon receipt of such notice, the Supplier shall, with all reasonable speed, replace the sub standard equipments, without cost to the Purchaser.

  - If the Supplier, having been notified, fails to remedy the defect(s) within seven days, the Purchaser may proceed to take such remedial actions as may be necessary, at the Supplier’s risk and expense and will have right to impose penalty without prejudice to any other rights which the Purchaser may have against the Supplier under the Contract.

  - The stores supplies shall be strictly in accordance with the Specifications / Standards and shall be of the best quality. The stores are demanded to carry the Supplier’s own guarantee of the items by the consignee.

  - If at any time during/after the supply if equipment is not found as per specification, sub standard or refurbished the bidder shall replace defective equipment at his own cost, immediately, failing which the total amount is recoverable from him and he will be black listed.

  - UPTIME GUARANTEE: The firm should provide uptime guarantee of 95%.

  - Downtime penalty Clause:

During the Guarantee / Warranty period, desired uptime of 95% of 365 days (24 hours), if downtime exceeds 5%, penalty in the form of extended warranty, double the number of days for which the equipment goes out of service will be applied. The vendor must undertake to supply all spares for optimal upkeep of the equipments for at least 10 years after handing over the unit to the Institute. If accessories /other attachment of the system are procured from the third party , then the vendor must produce cost of accessory/other attachment and the AMC / CMC  from the third party separately along with the main offer and the third party will have to sign the AMC / CMC with the Purchaser if required. In no case instrument should remain in non – working condition for more than 5 days, beyond which a penalty of 2 % of machine cost will be charged per day. The Principals or their agents are required to submit a certificate that they have satisfactory service arrangements and fully trained staff available to support the uptime guarantee.

  - Guarantee / Warranty period: The tenderers must quote for 5 years warranty from the date of completion of the satisfactory installation. The Warranty charges shall not be quoted separately otherwise the offer shall be summarily rejected. Also the Bidders should submit their quote for subsequent 5 years AMC (without spare parts) / CMC (include free labour, repair, other services & spare parts). Failure to comply this condition will entail the rejection of the Bids. The price comparison shall be made taking into account on basic price and post warranty AMC / CMC. The Rate Contracting Authority reserves the right to award AMC / CMC. A.M.C. (without spare parts) shall be quoted for equipments costing upto Rs.5.00 Lacs and C.M.C. (include free labour, repair, other services & spare parts) shall be quoted for equipments costing above Rs.5.00 Lacs. So the price of AMC / CMC should be quoted according to the cost of equipment.

  - SPARE PARTS: The separate price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee / warrantee period must be attached / enclosed along with the sealed quotation without this bid will not be considered. If any spares & accessories other than the price list attached/enclosed by the firm are required for future repair it will be borne by the firm only. The tenderers are required to furnish the list of spares along with their cost in the Financial Bid failing which their bids are liable to be rejected. The spare parts should be of standard quality. The bidder must take guarantee of availability of supply of spare parts upto 5 years beyond the warranty / gurantee period and must submit affidavit as per GCC clause 4.13.

  - TRAINING: Training of equipments within the stipulated time should be done by the suppiler at his cost. The time & place of training shall be stipulated by purchaser. Training should be of 2 doctors and 2 technicians of user department.

  - The Tenderers should clearly indicate the name of the Manufacturers / Beneficiary of the Letter of Credit, country of origin, place of shipment / air freightment etc.

  - Letter of Credit upto 75 % of order value shall be opened after receiving a bank guarantee for the entire 100%  of order value. The bank guarantee must cover the entire period up to the stage of handing over on turnkey basis.

  - Local agents quoting on behalf of their foreign suppliers must attach authority letter in their favour.

  - Successful tenderers will have to furnish performance Bank Guarantee for 10% contract value from any Nationalized Bank valid for the warranty period.

  - The rates quoted for the Stores/Equipments, under the reference, by the supplier shall in no event exceed the lowest price at which the suppliers of the Stores / Equipments of identical description are made to any other person / organization / institution during the period and should attach an undertaking (duly notarized).

  - Equipment should be brand new & of latest technology along with digital technique wherever applicable.

  - The Rate Contracting Authority reserves the right to increase the accessories and their numbers, payment will be made only for ordered accessories.

14. 	Payment

  - The method and conditions of payment to be made to the Supplier under the contract shall be specified in the SCC.

  - The Supplier’s request(s) for payment shall be made to the Purchaser in writing accompanied by an invoice describing, as appropriate, the Goods delivered and the service performed, and by documents, submitted pursuant to GCC Clause 10, and upon fulfillment of other obligations stipulated in the contract.

  - Payments shall be made by the Purchaser after submission of the claim by the Supplier. All sincere efforts will be made for payment of due amount which has been submitted to the purchaser within 30 days unless the situation being out of control of the purchaser. Proforma invoice should also be submitted.

  - Payment shall be made in Indian Rupees.

  - The payment of the claim / bill will be made after deduction of VAT as per rules of M.P. Commercial Tax Act Section 34 and other taxes from the bill.

  - No payment shall be made for rejected Stores. Rejected items must be removed by the supplier within two weeks of the date of rejection at their own cost and replace immediately. In case these are not removed these will be auctioned at the risk and responsibility of the suppliers without any notice.

  - Supply of equipments means – installation and commissioning and also test running at site. No separate charges will be paid separately on this account.

  - Payment will be made after installation, commissioning and successful test running at the site, due verification and subsequent satisfactory report of the user department.

15.	Prices

  - Prices charged by the Supplier for Goods delivered and Services performed under the Contract shall not vary from the prices quoted by the Supplier in its bid.

  - Recurring expenditure of the machine / equipment should be mentioned.

16.	Change orders

  - The Rate Contracting Authority may at any time, by written order given to the Supplier pursuant to GCC Clause 29 make changes within the general scope of the Contract in any one or more of the following:

        - the method of shipping or packing, installation;

        - Any other terms & conditions in public interest.

  - If any such change causes an increase or decrease in the cost of, or the time required, for the Supplier’s performance of any provision under the Contract, and equitable adjustment shall be made in the Contract Price or delivery schedule or both and the Contract shall accordingly be amended. Any claims by the Supplier for adjustment under this clause must be asserted within fifteen (15) days from the date of the Supplier’s receipt of the Purchaser’s change order.

  - The Purchase Orders on approved rates will be placed by the Purchaser.

17. 	Contract Amendments.

  - Subject to GCC Clause 16, no variation in or modification of the terms of the Contract shall be made except by written amendment signed by both the parties.

18.	Assignment

  - The Supplier shall not assign, in whole or in part, its obligations to perform under the Contract, except with the Purchaser’s prior written consent.

19.	Delays in the Supplier’s Performance

  - Delivery of the Goods and performance of the Services shall be made by the Supplier in accordance with the time schedule specified by the Purchaser in the Supply order.

  - If at any time during performance of the Contract, the Supplier should encounter conditions impeding timely delivery of the Goods and performance of the Service, the Supplier shall promptly notify the Purchaser in writing of the fact of the delay, its likely duration and its cause(s). As soon as practicable after receipt of the Supplier’s notice the Purchaser shall evaluate the situation and may at its discretion extend the supplier’s time for performance.

  - Except as provided under GCC Clause 22, a delay by the Supplier in the performance of its delivery obligations shall render the Supplier liable to the imposition of liquidated damages pursuant to GCC Clause 20, unless an extension of time is agreed upon pursuant to GCC Clause 19.2 without the application of liquidated damages.

20.	Liquidated Damages

  - Subject to GCC Clause 22, if the Supplier fails to deliver any or all the Goods or to perform the services within the period(s) specified in the supply order, the Purchaser shall, without prejudice to its other remedies under the Contract, deduct from the Contract Price, as liquidated damages, a sum equivalent to the percentage specified in SCC of the delivered price of the delayed goods or unperformed services for each week or part thereof of delay until actual delivery or performance, up to a maximum deduction of percentage specified in SCC. Once the maximum is reached, the Purchaser may consider termination of the Contract pursuant to GCC Clause 21.

21.	Termination for Default

Contract may be terminated by the Rate Contract Authority if:

  - If the supplier fails to execute the supply within the stipulated time, the Purchaser is at liberty to make alternative purchase, in the event of making ALTERNATIVE PURCHASE, the supplier will be imposed penalty apart from the forfeiture of Performance Guarantee. The excess expenditure over and above contracted prices incurred by the Purchaser in making such purchases from any other sources or in the open market or from any other supplier who has quoted higher rates and other losses sustained in the process, shall be recovered from the Performance Security or from any other money due and become due to the Supplier and in the event of such amount being insufficient, the balance will be recovered personally from the Supplier. The penalty would be as under:

      - First extension 31st day thereof from the date of issue of supply order – 3% of supplied ordered item.

      - Second & maximum after 45 days from the date of issue of supply order – 5% of supplied ordered item.

      - The order will be deemed cancelled after expiry of 60 days from the issue date.

  - The order may be cancelled after expiry of delivery period as mentioned in the supply order and the supplier shall also suffer forfeiture of the Performance Security and shall invite other penal action like blacklisting / disqualification from participating in present and future tenders.

  - Rate Contracting Authority will be at liberty to terminate by assigning justifiable reason thereof the contract either wholly or in part on one month notice. The Supplier will not be entitled for any compensation whatsoever in respect of such termination.

  - If the Supplier, in the judgment of the Rate Contracting Authority has engaged in corrupt or fraudulent practices in competing for or in executing the contract.

For the purpose of this Clause.

“Corrupt practice” means offering, giving, receiving or soliciting of anything of value to influence the action of a public official in the procurement process or in contract execution.

“Fraudulent practice” means a mis-presentation / hiding of facts in order to influence a procurement process or the execution of a contract to the detriment of the other bidders, and includes collusive practice among Bidders (prior to or after bid submission) designed to establish bid prices at artificial noncompetitive levels and to deprive the other bidders of the benefits of free and open competition.

  - For infringement of the stipulations of the contract or for other justifiable reasons, the contract may be terminated by the Rate Contracting Authority, and the supplier shall be liable for all losses sustained by the Rate Contracting Authority, in consequence of the termination which may be recovered personally from the supplier or from his properties, as per rules.

  - Non performance of any of the contract provisions will disqualify a firm to participate in the tender for the next five years.

  - In all the above conditions, the decision of the Rate Contracting Authority shall be final and binding.

22.	Force Majeure

  - Not with standing the provision of GCC Clause 19, 20, 21, the Supplier shall not be liable for forfeiture of its performance security, liquidated damages, penalty or termination for default, if and to the extent that, its delay in performance or other failure to perform its obligations under the Contract is the result of an event of Force Majeure.

  - For purpose of this Clause, “Force Majeure” means an event beyond the control of the Supplier and not involving the Supplier’s fault or negligence and not foreseeable. Such events may include, but are not limited to, acts of the Rate Contracting Authority either in its sovereign or contractual capacity, wars or revolution, fires, floods, epidemics, quarantine restrictions and freight embargoes.

  - If a Force Majeure situation arises, the Supplier shall promptly notify the Rate Contracting Authority in writing with adequate proof of such conditions and the cause thereof. Unless otherwise directed by the Rate Contracting Authority in writing the Supplier continue to perform its obligations under the Contract as far as it is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by Force Majeure event.

23.	Termination for insolvency

  - The Rate Contracting Authority may at any time terminate the contract by giving written notice to the Supplier, if the Supplier becomes bankrupt or otherwise insolvent. In this event, termination will be without compensation to the Supplier, provided that such termination will not prejudice or affect any right of action or remedy which has accrued or will accrue thereafter to the Rate Contracting Authority.

24.	Termination for Convenience

  - The Rate Contracting Authority, by written notice sent to the Supplier, may terminate the Contract, in whole or in part, at any time for its convenience. The notice of termination shall specify that termination is for the Rate Contracting Authority’s convenience, the extent to which performance of the Supplier under the contract is terminated, and the date upon which such termination become effective.

  - The Goods that are complete and ready for shipment within 30 days after the Supplier’s receipt of notice of termination shall be accepted by the Rate Contracting Authority at the Contract terms and prices. For the remaining Goods, the Rate Contracting Authority may elect:

          - to have any portion completed and delivered at the Contract terms and prices; and / or

          - to cancel the remainder and pay to the Supplier an agreed amount for partially completed Goods and for materials and parts previously procured by the Supplier.

25. 	Resolution of Disputes

  - The Rate Contracting Authority and the Supplier for the rate contracts & purchaser and supplier for supply order, supply, delivery and payment and other issues shall make every effort to resolve amicably by direct informal negotiations any disagreement or dispute arising between them under or in connection with the Contract.

  - If, after thirty (30) days from the commencement of such informal negotiations, the Rate Contracting Authority and the Supplier & purchaser and the supplier have been unable to resolve, amicably a Contract dispute, either party may require that the dispute be referred for resolution to the formal mechanisms specified in the SCC. These mechanisms may include, but are not limited to, conciliation mediated by a third party, adjudication in an agreed national or international forum, and/or international arbitration.

i. 	Any dispute or difference in respect of which a notice of intention to commence arbitration has been given in accordance with this clause shall be finally settled by arbitration. Arbitration may be commenced prior to or after delivery of the goods under the contract.

ii.	Arbitration proceedings shall be conducted in accordance with the rules of procedure specified in SCC.

  - Notwithstanding any reference to arbitration herein the parties shall continue to perform their respective obligations under the contract unless they otherwise agree.

26.	Limitation of Liability

  - Except in cases of criminal negligence or willful misconduct, and in the case of infringement pursuant to Clause 6.

i. 	the supplier shall not be liable to the Rate Contracting Authority, whether in contract, tort, or otherwise, for any indirect or consequential clause or damage, loss of use, loss of production or loss of profits or interest costs, provided that this exclusion shall not apply to any obligation of the suppliers to pay liquidated damages to the Rate Contracting Authority, and

    - the aggregate liability of the supplier to the Rate Contracting Authority, whether under the contract, in tort or otherwise, shall not exceed the total ordered price, provided that this limitations shall not apply to the cost of replacing sub-standard/defective goods.

27.	Governing Language

  - The contract shall be written in English language. Subject to GCC Clause 28, English language version of the Contract shall govern its interpretation. All correspondence and other documents pertaining to the Contract which are exchanged by the parties shall be written in the same language.

28.	Applicable Law

  - The Contract shall be interpreted in accordance with the laws of the Union of India.

29.	Notices

  - Any notices given by one party to the other, pursuant to this Contract, shall be sent to other party in writing, confirmed in writing to the other Party’s address specified in SCC.

  - A notice shall be effective when delivered or on the notice’s effective date, whichever is later.

30.	Taxes and Duties

  - In case of any enhancement in Excise Duty due to notification of the Government after the date of submission of tenders and during the tender period, the quantum of additional excise duty so levied will be paid extra, if the rates of excise duty prevailing at the time of tender, has been shown extra and actually paid by the supplier. For claiming the additional cost on account of the increase in Excise Duty, the supplier should produce a letter from the concerned Excise authorities for having paid additional Excise Duty on the goods supplied to Purchaser and also must claim the same in the invoice separately.

  - Suppliers shall be entirely responsible for all taxes, duties license fees, octroi, road permits, etc. incurred until delivery of the contracted Goods to the Purchaser. However, Sales tax/VAT (not surcharge in lieu of Sales Tax/VAT) in respect of the transaction between the Purchaser and the Supplier shall be payable extra, if so stipulated in the supply order.

31.	Fall Clause

  - Prices charged for supplies under Rate Contract by the supplier should in no event exceed the lowest prices at which he offers to sell or sells the stores of identical description to any other State Government / DGS & D/ Public Undertaking during the period of the contract.

  - If at any time during the period of contract, the price of tendered items is reduced or brought down by any law or Act of the Central of State government, the supplier shall be bound to inform Rate Contracting Authority immediately about such reduction in the contracted prices, in case the supplier fails to notify or fails to agree for such reduction of rates, the Rate Contracting Authority will revise the rates on lower side. If there is a price increase for any product after quoting the rates, the bidder will have to supply the item as per quoted rates. This office will not accept any higher rates after wards.

  - If at any time during the period of contract, the supplier quotes the sale price of such Equipments or sells such Equipments to any other State Govt. / DGS&D and Public Undertakings at a price lower than the price chargeable under the rate contract he shall forthwith notify such reduction to Rate Contracting Authority and the prices payable under the rate contract for the Equipments supplied  from the date of coming into force of such price stands correspondingly reduced as per above stipulation however reduction shall not apply to :-

(a)  	Export by the supplier

(b)  	For all contracts entered into prior to the date of the tender or for any backlog of pending orders.

  - Within six months of the commencement of the rate contract and at the rate contract period a certificate in the following forms will have to be submitted by the supplier :-

I/We certify that the stores of description identical to the store supplied to the Govt. of M.P. under the contract herein have not been sold by me/us to any other State Govt. / Central Govt. / DGS & D / Public Undertaking during the period of the rate contract of Madhya Pradesh under the contract / except for the quantity of under sub-clause (a) & (b) of the clause 31.3.

32.	Jurisdiction

  - In respect of all disputes or claims related with Rate Contracts out of or under this contract, Indore Court alone shall have jurisdiction to entertain the same.

  - In respect of all disputes or claims related with Supply, Payments and any other out of or under this contract, the concerned Court of Purchaser’s place shall have jurisdiction to entertain the same.

CHAIRMAN, PURCHASE COMMITTEE &

DEAN, M.G.M MEDICAL COLLEGE,    A.B. ROAD INDORE (M.P.)

SECTION IV :

SPECIAL CONDITIONS OF CONTRACT (SCC)

SECTION IV :

SPECIAL CONDITIONS OF CONTRACT (SCC)

TABLE OF CLAUSES

| Clause No. | Topic | Page No. |
| --- | --- | --- |
| 1 | Definitions (GCC Clause 1) | 47 |
| 2 | Performance security (GCC Clause 7) | 47 |
| 3 | Inspection and tests (GCC Clause 8) | 47 |
| 4 | Annual / Comprehensive Maintenance Contract) & Training | 49 |
| 5 | Packing (GCC Clause 9) | 49 |
| 6 | Delivery and documents (GCC Clause 10) | 49 |
| 7 | Insurance (GCC Clause 11) | 50 |
| 8 | Payments (GCC Clause 14) | 50 |
| 9 | Prices (GCC Clause 15) | 50 |
| 10 | Liquidated damages & deduction in payment (GCC Clause 20) | 51 |
| 11 | Resolution of disputes (GCC Clause 25) | 51 |
| 12 | Notices (GCC Clause 29) | 52 |
| 13 | Supplier Integrity | 52 |
| 14 | Supplier’s obligations | 52 |
| 15 | Patent right (GCC Clause 6) | 53 |
| 16 | Progress of Supply | 53 |
| 17 | Section V: Bid Form | 54 |

Special Conditions of Contract

The following special conditions of contract shall supplements the general conditions of contract whenever there is a conflict, the provisions herein shall prevail, over those in the general conditions of contract the corresponding clause numbers of the general conditions is indicated in parentheses.

1. 	Definitions (GCC Clause 1)

GCC 1.1 (g)    (a) 	The Purchaser is concerned Deans / Principal / Superintendents of various Govt. / Autonomous Medical Colleges, Dental College and Associated Hospitals of Madhya Pradesh which is also Good’s Receiving Authority.

GCC 1.1 (i)     (b) 	The Supplier is the individual or firm supplying the Goods and Services under this Contract.

GCC 1.1 (h)	(c) 	The Purchaser Country is India.

GCC 1.1 (l)	(d) 	The project site is as per supply order.

2.	Performance security (GCC Clause 7)

2.1	The supplier shall be required to pay 10% performance security of the order value or maximum Rs. 15 lac which ever is less. The performance security should be paid upfront in respect of each supply order or before the due date fixed by the Purchaser, valid up to the end of guarantee / warranty period for performance obligations including warranty obligations.

  - Substitute clause 7.4 of the GCC by the following.

The performance security will be discharged by the Purchaser and returned to the supplier not later than 60 days following the date of completion of the supplier’s satisfactory performance obligations including the warranty obligations under the contract.

  - Add as clause 7.5 to the GCC the following:-

In the event of any contract amendment, the supplier shall, within 07 days of receipt of such amendment furnish the amendment to the performance security, rendering the same valid for the duration of the contract as amended for further period of 60 days thereafter.

3.	Inspection and tests (GCC Clause 8)

If purchaser wishes:

- The inspections shall be carried out by the appointed Technical Committee or Inspection Agency at the premises of the suppliers / godown or stores of the supplier / at point of delivery / installation. Inspection and testing charges for the above purpose shall be borne by the supplier.

- Inspection note will be issued by the inspection committee verifying the specification, performance, details of accessories supplied with the machine, test certificate issued by the respective authority etc. as decided by the purchasing committee.

- The machine will be dispatched only after the inspection procedure has been followed and inspection note issued to accept the consignment.

- The consignee may also draw the sample, at random, from the consignment within 45 days of their receipts, and get them re-tested to satisfy whether the lots conform to the laid down specification. In the event of the sample failing to conform to specification, the consignee shall reject the batch of supply and inform the supplier for arranging replacement of the rejected batches at his own cost.

- When the inspection conducted on the premises of the supplier, all reasonable facilities and assistance including access to drawing and production data shall be furnished to the inspectors at no charge to the Purchaser.

- In the event of the sample of EQUIPMENTS failing quality test and found to be not as per specification the Purchaser is at liberty to make alternative purchase of the items, of EQUIPMENTS for which the supply orders have been placed, from any other sources or in the open market or from any other suppliers who might have quoted higher rate at Bid and the cost of the supplier and in such cases the Purchaser has every right to recover the excess cost from supplier’s performance security.

- If any items of equipments supplied by the supplier have been partially or wholly used or consumed after supply and are subsequently found to be in bad order, unsound, inferior in quality or description or otherwise faulty or unfit for consumption and if payment had already been made to him then the contract price or prices of such articles or things will be recovered from the supplier,. The supplier will not be entitled to any payment, whatsoever, for items of equipments found to be NOT OF STANDARD QUALITY whether consumed or not and the purchaser is entitled to deduct the cost of such equipments from any amount payable to the supplier. On the basis of nature of failure, the product / supplier will be moved for black listing.

- For equipments labelled as NOT OF STANDARD QUALITY, the concerned administration will be informed for initiating necessary action against the supplier and that product shall be banned / black listed and no further supplies will be accepted from him till he is legally discharged. The supplier shall also not be eligible to participate in tenders for supply of such equipments for a period of five subsequent years.

4.	Annual (without spare parts) (AMC) / Comprehensive (include free labour, repair, other services & spare parts) Maintenance Contract (CMC) & Training

  - The Bidder shall also quote charges for Annual (without spare parts) / Comprehensive (include free labour, repair, other services & spare parts) Maintenance Contract for the next five years after the expiry of five years warranty period in Annexure-XII.

  - The bidder shall provide operational training to Technician staff / operator for minimum of 3 days by the expert or as instructed at the time of agreement.

  - The bidder should take guarantee of the availability of all spare parts for a minimum period of 10 years from the date of installation.

  - Genuine equipments and instruments etc. should be supplied. Tenderers should indicate the source of supply i.e. name and address of the manufacturers from whom the items are to be imported.

5.	Packing (GCC Clause 9)

Add as clause 9.3 of the GCC of the following:-

Packing Instructions: The Supplier will be required to make separate packages for each Consignee. Each package will be marked on three sides with proper paint/indelible ink, the following:

(i) Project	(ii) Contract No. (iii) Country of Origin of Goods (iv) Supplier’s Name; and (v) Packing list reference number.

  - Packing should be able to prevent damage or deterioration during transit.

  - In the event of items of equipments supplied found to be not as per specifications in respect of their packing, the Purchaser is at liberty to make alternative purchase of the items of equipments for which the supply orders have been placed from any other sources or in the open market or from any other bidder who might have quoted higher rates at the risk and the cost of the supplier and in such cases the Purchaser has every right to recover the cost and imposes penalty as mentioned in GCC clause 21.1.

6.	Delivery and documents (GCC Clause 10)

Upon delivery of the goods, the supplier shall submit the following documents to the Purchaser.

- Three copies of the supplier invoice showing Goods description, quantity, unit price, and total amount.

- Acknowledgement of receipt of goods from the consignee(s).

- Installation certificate signed by respective consignee.

- Manufacturer’s / supplier’s warranty certificate.

- Inspection certificate issued by the nominated inspection agency, and the Supplier’s factory inspection report; and

- Certificate of origin.

- Photocopy of all test report of all equipments etc. should be submitted with every delivery challan.

7.	Insurance (GCC Clause 11)

For delivery of goods at site, the insurance shall be obtained by the supplier in an amount equal to the value of the goods from final destinations as specified in the supply order of “All Risks” basis including war Risks and strike.

Should any loss or damage occurs, the supplier shall:

- Initiate and pursue claim till settlement, and

- Promptly make arrangement for replacement of any damaged item/s irrespective of settlement of claim by the underwriters.

8.	Payments (GCC Clause 14)

Payment for goods and services shall be made in Indian Rupees as follows:-

  - No advance payments towards cost of equipments etc. will be made to the supplier.

  - All payments shall be made by way of crossed cheques drawn in favour of the supplier.

  - All bills / invoices should be raised in triplicate in the name of Concerning Purchaser.

  - Payment will be made after completion of supply of goods / service on turn key basis as per supply order, installation, commissioning and successful test running at the site, due verification and subsequent satisfactory report of the user department. Payments shall be made by the Purchaser after submission of the claim by the Supplier. All sincere efforts will be made for payment of due amount which has been submitted to the purchaser within 30 days unless the situation being out of control of / unforeseen for the purchaser. Proforma invoice should also be submitted.

  - FALL CLAUSE: If, at any time, during the said period, the supplier reduce the said prices of such Stores/ Equipment or sales such stores to any other person/organization at a price lower than the chargeable, he shall forthwith notify such reduction or sale to the PURCHASER and the price payable for the Stores supplied after the date of coming into force of such reduction or sale shall stand correspondingly reduced.

9.	Prices (GCC Clause 15)

Substitute clause 15.1 of the GCC with the following:

Prices payable to the supplier as stated in the contract shall not be subject to adjustment during performance of the contract

10.	Liquidated damages & deduction in payment (GCC Clause 20)

  - For delay :

Substitute GCC clause 20.1 by the following:

Subject to GCC clause 20, if the supplier fails to deliver any or all the goods or perform the services within the time period(s) specified in the contract. The Purchaser shall, without prejudice to its other remedies under the contract, deduct from the contract price as liquidated damages, as shown below of the delivered price of the delayed goods or unperformed services for each week of delay or part thereof until actual delivery or performance up to maximum deduction of 5% of the delayed goods or services contract price. Once the maximum is reached, the purchaser may consider termination of the contract.

      - First penalty: 31st day from the date of issue of supply order – 3% of supply ordered item.

      - Second penalty: After additional 45 days from the date of issue of supply order – 5% of supply ordered item.

      - The order will be deemed cancelled after expiry of 60 days from the issue date

  - Purchaser has every right to receive supply even after expiry of delivery period as mentioned in the supply order and in such case, liquidated damages will be levied @ 3% of the delivery price of the delayed goods or unperformed services for each week of delay or part thereof until actual delivery or performance.

  - Supply in damaged condition shall not be accepted. In case of damage in the packing, the supply will be accepted only after levying penalty or replacement of damaged supply on the total value of supply to that particular / other designated place.

  - Supply must be in toto i.e. not in fraction.

11.	Resolution of disputes (GCC Clause 25)

Add as GCC clauses 25.4 and 25.5 the following:

  - The dispute resolution mechanism to be applied pursuant to GCC clause 25 shall be as follows:

- In case of dispute or difference arising between the Rate Contracting Authority / Purchaser and supplier relating to any matter arising out of or connected with this agreement, such disputes or difference shall be settled in accordance with the Arbitration and Conciliation Act, 1996. The Next Higher Authority shall be the Arbitrator.

  - The Venue of Arbitration shall be at concerned place of next higher authority of R.C.A. / Purchaser.

12.	Notices (GCC Clause 29)

For the purpose of all notices, the following shall be the address of the Rate Contracting Authority & Purchaser and supplier:

Rate Contracting Authority:	The Chairman, Purchase Committee &

Dean, M.G.M. Medical College,

Indore (M.P.)

Purchaser	:

        - Dean, M.G.M. Medical College, A.B.Road, Indore

        - Dean, Gandhi Medical College, Bhopal

        - Dean, Gajra Raja Medical College, Gwalior

        - Dean, Shyam Shah Medical College, Rewa

        - Dean, N.S.C.B. Medical College, Jabalpur

        - Dean, Bundelkhand Medical College, Sagar

        - Principal, Dental College, A.B. Road, Indore

        - Jt. Director cum Superintendent, M.Y. Hospital, MYH Road, Indore

        - Superintendent, Govt. Cancer Hospital, Near MY Hospital, Indore

        - Superintendent, Chacha Nehru Bal Chikitsalaya Avum Anusandhan Kendra, Behind MY Hospital, Indore

        - Superintendent, Mental Hospital, Banganga, Indore

        - Superintendent, Hamidia Hospital, Bhopal

        - Superintendent, Sultania Zanana Hospital, Bhopal

        - Superintendent, Jayarogya Hospital, Gwalior

        - Superintendent, Mental Hospital, Gwalior

        - Superintendent, S.G.M. Hospital, Rewa

        - Superintendent, N.S.C.B. Medical College Hospital, Jabalpur

        - Superintendent, Govt. Medical College Hospital, Sagar

Supplier 	:	(To be filled at the time of Contract Signature)

---------------------------------------------------------

13.	Supplier Integrity

The supplier is responsible for and obliged to conduct all contracted activities in accordance with the contract using state-of-the-art methods and economic principles and exercise all means available to achieve the performance as specified in the contract.

14.	Supplier’s obligations

The supplier is obliged to work closely with the R.C.A. & Purchasers staff, act within its own authority and abide by directives issued by the Purchaser and implementation activities.

The supplier will abide by the job safety measures prevalent in India and will free the purchase from all demands or responsibilities arising from accidents or loss of life the cause of which is the supplier’s negligence. The supplier will pay all indemnities arising from such incidents and will not hold the Purchaser responsible or obligated.

The supplier is fully responsible for managing the activities of its personnel or sub contracted personnel and will hold itself responsible for any misdemeanors.

The Supplier will treat all data and information about the Rate Contracting Authority / Purchaser, obtained in the execution of his responsibilities, in strict confidence and will not reveal such information to any other party without the prior written approval of the Rate Contracting Authority / Purchaser.

15.	Patent right (GCC Clause 6)

In the event of any claim asserted by a third party of infringement of copyright , patent, trademark or industrial design rights arising from the use of goods or any part thereof in the Purchaser’s country, the supplier shall act expeditiously to extinguish such claim. If the supplier fails to comply and the Purchaser is required to pay compensation to a third party resulting from such infringement, the supplier shall be responsible for the compensation including all expenses court cost and lawyers fees. The Purchaser will give notice to the supplier of such claim, if it is made, without delay.

16.	Progress of Supply

Supplier 	:	(To be filled at the time of Contract Signature)

---------------------------------------------------------

---------------------------------------------------------

Supplier shall regularly at every 7 days interval shall intimate item wise progress of supply in writing, to the Purchaser as under:

  - Quantity offered for inspection and date :

  - Quantity accepted / rejected by inspecting agency and date:

  - Quantity dispatched / delivered to consignee and date :

  - Quantity where incidental services have been satisfactorily completed with date :

  - Quantity where rectification / replacement effected / completed on receipt of any communication from consignee / Purchaser with date :

(In case of state-wise inspection, details required may also be specified).

CHAIRMAN, PURCHASE COMMITTEE &

DEAN, M.G.M MEDICAL COLLEGE,

A.B. ROAD INDORE (M.P.)

SECTION V

Ref. Clause No. 17.1(p) of ITB

BID FORM

Date : 17.04.2013

Tender No. CPC/EQP-MOT/13-14/03

To,

Chairman, Purchase Committee &

Dean, M.G.M. Medical College

A.B.Road, Indore (M.P.)

I/We, the undersigned, declare that:

    - I/We have examined the bidding documents including Addenda Nos. ……. (insert numbers), the receipt which is hereby acknowledged.

    - I/We have gone through all terms and conditions of the tender document before submitting the same. I/We hereby agree to all terms and conditions as stipulated in the tender document and offer to supply and deliver ………………….. (Brief description of equipments) in conformity with the bidding documents in accordance with the schedule of prices attached herewith and made part of this bid.

    - I/We undertake, if our bid is accepted, to deliver the goods in accordance with delivery period specified in the supply order.

    - I/We agree to abide by this bid for a period of 180 (numbers) days after the date fixed for bid opening and shall remain binding upon us and may be accepted at any time before the expiration of that date.

    - If our bid is accepted, we commit to deposit  a performance security in accordance with GCC clause 7 & SCC clause 2 for the due performance of the contract.

    - Until a formal contract is prepared and executed, this bid together with your written acceptance thereof and your notification of rate contract shall constitute a binding contract between us.

    - I/We undertake if at any time, it is found that any information furnished by us to the Rate Contracting Authority, either in our bid or otherwise, is false, the Rate Contracting Authority servers the right to terminate the contract without assigning any reasons, forfeiting the bid security or performance security and blacklisting us for a period of 5 years.

    - I/We understand that you are not bound to accept the lowest or any bid you may receive.

    - I/We hereby submit our tender for the  ______________________________.

    - I/We now enclosing herewith the E.M.D. No.___________ dated _________.

    - I/We have noted that overwritten entries shall be deleted unless duly cut & re-written and initialed.

    - Tenders are duly signed (No thumb impression should be affixed).

    - I/We undertake to sign the contract / agreement, if required, within 15 (fifteen) days from the date of issue of the letter of acceptance, failing which our/my security money deposited may be forfeited and our/my name may be removed from the list of suppliers.

Dated this ………………….. day of ……………….. 2013.

(Signature) ……………..

(in the capacity of :……………….)

Duly authorized to sign for and on behalf of ………………..

Witness 1

Witness 2

TABLE OF ANNEXURES

| Sr. No. | Topic | Page No. |
| --- | --- | --- |
| 1 | Annexure – I (Sales Tax Clearance Certificate) | 56 |
| 2 | Annexure-II (Manufacture Authorization Form) | 58 |
| 3 | Annexure – III (Declaration / Undertaking Form) | 59 |
| 4 | Annexure – IV (Proforma for Performance Statement) | 60 |
| 5 | Annexure – V (Annual Turnover Statement) | 61 |
| 6 | Annexure – VI (Specification of require Equipments) | 62 |
| 7 | Annexure – VII (Performance Security Form) | 63 |
| 8 | Annexure – VIII (Contract Agreement) | 64 |
| 9 | Annexure – IX (Details of Manufacturing Unit) | 67 |
| 10 | Annexure – X (The Price Schedule) | 68 |
| 11 | Annexure – XI (Price Schedule for AMC / CMC) | 69 |
| 12 | Annexure – XII (Details of Service Centre in M.P.) | 70 |
| 13 | Check List | 71 |
| 14 | Requirement of Equipments (Departmentwise) | 73 |

ANNEXURE – I

Ref. Clause No. 17.1 (g) of ITB

FORM OF CERTIFICATE OF SALES TAX / VAT VERIFICATION TO BE PRODUCED BY AN APPLICANT FROM THE CONTRACT OR OTHER PATRONAGE AT THE DISPOSAL OF THE GOVERNMENT OF MADHYA PRADESH

(To be filled up by the applicant)

- Name of style in which the applicant is addressed or assessable to sales tax / VAT addresses or assessment.

- a. 	Name and address of all companies , firms or associations or   persons in which the applicant is interested in his individual or fiduciary capacity

- Places of business of the applicant (all places of business should be mentioned)

- The Districts, blocks and division in which the applicant is assessed to sales tax / VAT (all places of business should be furnished)

- a. 	Total contract amount or value of patronage received in the preceding three years

2009-10

2010-11

2011-12

- Particular of sales – Tax / VAT for the preceding three years

| Year | Total T.O. (Turnover) be assessed (Rs) | Total Tax assessed (Rs) | Total Tax Paid (Rs) | Balance due (Rs) | Reasons for Balance (Rs) |
| --- | --- | --- | --- | --- | --- |
| 2009-10 |  |  |  |  |  |
| 2010-11 |  |  |  |  |  |
| 2011-12 |  |  |  |  |  |

- If there has been no assessment in any year, whether any returns were submitted? if yes,  the division in which the returns were sent?

- Whether any penal action or proceeding for the recovery of Sales tax / VAT is pending?

- The name and address of Branches, if any :

I declare that that the above information is correct and complete to the best of my knowledge and belief.

Signature of Applicant:

Address:

Date:

(To be filled up by the Assessing Authority)

In my opinion, the applicant mentioned above has been / has not been / doing everything possible to pay the tax demands promptly and regularly and to facilitate the completion of pending proceeding.

Date Seal	:	Deputy / Asstt. Commercial Tax – Officer

Deputy Asstt.

Note: A separate certificate should be obtained in respect of each of the place of business of the applicant from the deputy commercial tax officer or Assistant commercial tax officer having jurisdiction over that place.

ANNEXURE – II

Ref. Clause No. 17.1 (q) of ITB

MANUFACTURER’S AUTHORIZATION LETTER

No………………………………… Dated………………

To,

Dear Sir,

Tender No.:

We ___________________ an established and reputable Manufacturers of ___________________ having factories at ___________ and ________ do hereby agree to supply ________________ confirming to the required specification and required quantity to M/s ____________________ (Bidder) as offered by them to supply against the above stated Tender. This is also certified that M/s _____________________________ is our authorized distributor / importer since ____________ (month & year should filled), and his performance is satisfactory.

We hereby extend our full guarantee and warranty as per Clause 15 of the General Conditions of Contract for the supply against this invitation for Bid by the above firm.

Yours faithfully,

(name)

for and on behalf of M/s ____________________________ (Name of manufacturers)

Note: This letter should be signed by a person competent and having authority to sign on behalf of manufacturer, and should be duly Notarized.

ANNEXURE – III

Ref. Clause No. 17.1 (T) of ITB

DECLARATION / UNDERTAKING

I/We/ M/s.___________________________ represented by its Proprietor / Managing Partner / Managing Director having its Registered Office at _______________________ and its Factory Premises at ____________ do declare that I/We have carefully read all the conditions of tender in Ref. No._____________  for supply of equipment, floated by the Purchase Committee, and accept all conditions of Tender.

I/We agree that the Purchaser has rights of forfeiting the Bid Security and or Performance Security Deposit and blacklisting me/us for a period of 7 years if any information furnished by us proved to be false at the time of inspection and not complying to the tender conditions.

Signature of the Bidder

Name & Address in capital letters with Designation

To be duly Notarized.

ANNEXURE – IV

Ref. Clause No. 17.1(e) of ITB

PROFORMA FOR ITEMWISE LIST OF INSTALLATIONS IN LAST THREE YEARS OF THE MANUFACTURER’S

Name of the Manufacturer __________________________________________________

________________________________________________________________________

Sl.  Name of installed machines    Name of the Purchaser & address    Date of            Quantity

No.    and model                          with phone number                         installation

1                                                      2                                  3                      4

1

2

3

4

5

6

7

8

9

10

11

12

Signature and seal of the Bidder

Annexure-V

Ref. clause No. 13.3(l) of ITB

ANNUAL TURNOVER STATEMENT

The annual Turnover of M/s ……………………………………………………. For the past three years are given below and certified that the statement is true and correct.

Turnover in Crore (Rs.)

| Sr No. | Year | Turnover in Crores (Rs) |
| --- | --- | --- |
| 1. | 2009-10 |  |
| 2. | 2010-11 |  |
| 3. | 2011-12 |  |

Date :

Seal:

Signature of Auditor ⁄

Chartered Accountant

(Name in Capital)

ANNEXURE-VI

Ref. Clause No. 11.1 of ITB

SPECIFICATIONS OF EQUIPMENTS

Tender No.

| Sr No | Item Code | Name of Item / Equipment | Specification Required by Purchaser | Make & Model | Specification as quoted by bidder | Compliance / Deviations |
| --- | --- | --- | --- | --- | --- | --- |

ANNEXURE-VII

Ref. Clause No. 7.3 of GCC

PERFORMANCE SECURITY FORM

To: ………………………………………………	(Name of Purchase)

Whereas ………………………………………...      (Name of Supplier)

hereinafter called “the supplier” has undertaken , in pursuance of Contract No. ……… dated…… 2013 to supply …………………………..  [description of goods and related services]  hereinafter called “the Contract”.

AND WHEREAS it has been stipulated by you in the said Contract that the Supplier shall furnish you with a Bank Guarantee by a recognized bank for the sum specified therein as security for compliance with the Suppliers performance obligations in accordance with the Contract.

AND WHEREAS we have agreed to give the Supplier a Guarantee:

THEREFORE, WE hereby affirm that we are Guarantors and responsible to you, on behalf of the Supplier, up to a total of _____________ (Amount of the Guarantee in Words and Figures) and we undertake to pay you, upon your first written demand declaring the Supplier to be in default under the Contract and without cavil or argument, any sum or sums within the limit of __________________ (amount of Guarantee) as aforesaid, without needing to prove or to show grounds or reasons for your demand or the sum specified therein.

This guarantee is valid until the ………………….. day of …………….. 2013   .

Signature and Seal of Guarantors

-------------------------------------------------

-------------------------------------------------

Date ………..………2013

Full Address of the Bank:

-------------------------------------------------

-------------------------------------------------

Telephone No. -------------------------------

Fax No. ---------------------------------------

Email Address -------------------------------

ANNEXURE-VIII

Ref. Clause No.31 of ITB

CONTRACT AGREEMENT FORM

(Tender No.______)

THIS CONTRACT AGREEMENT made the ………………………day of …….. 2013 between Rate Contracting Authority (Chairman, Purchase Committee and Dean, M.G.M. Medical College, Indore M.P.) (Name of Rate Contracting Authority) of India (country of Rate Contracting Authority) (hereinafter  called   “the Rate Contracting Authority” )   of    one  part  and                                          M/s ……………………………….  (name of supplier) of …………………………………. (city and country of supplier) (hereinafter called “the supplier”) of the other part :

WHEREAS the Rate Contracting Authority invited bids for certain goods and ancillary services viz. EQUIPMENTS (Brief description of goods” and services) and has accepted a bid by the supplier for the supply of those goods and services.

NOW THIS AGREEMENT WITNESSETH AS FOLLOWS:

- In this agreement words and expression shall have the same meaning as are respectively assigned to them in the conditions of contract referred to :

- The following documents shall constitute the contract between the Rate Contracting Authority and the supplier, and each shall be read and construed as an integral part of the contract :

  - This contract agreement :

  - Instructions of contract :

  - General conditions of contract :

  - Special conditions of contract :

  - Technical Specifications :

  - The supplier’s bid and original price schedules

  - The Rate Contracting Authority’s notification of rate contract.

- This contract shall prevail all other contract documents. In the event of any discrepancy or inconsistency with the contract documents, then documents shall prevail in the order listed above.

- In consideration of the payments to be made by the Purchaser to the supplier as hereinafter mentioned, the supplier hereby covenants with the Purchaser to provide the goods and services and to remedy defects therein in conformity in all respects with the provisions of the contract.

- The Purchaser hereby covenants to pay the supplier in consideration of the provision of the goods and services and the remedying of defects therein, the contract price or such as may become payable under the provisions of the contract at the times and in the manner prescribed by the contract.

Brief particulars of the goods and services which shall be supplied / provided by the supplier are as under:-

| Sr No. | Item Code | Item Description | Unit | F.O.R. Rate per unit (Rs.)* |
| --- | --- | --- | --- | --- |

* The above rates are inclusive of excise duty, transportation, insurance, inspection & testing charges and any incidental charges, but exclusive of CST/VAT.

- The prices shall be valid for one year from the date of agreement, unless revoked and thereafter for a further period as agreed upon mutually.

- The supplier shall agree to deposit inspection and testing charges and service tax as per tender conditions, in advance by cash / demand draft, against the value of supply order.

- The supplier shall agree to deposit 10% performance security, along with as mentioned at point no. 7 (above), in advance by FDR / Bank Guarantee, against the value of particular supply order for a period of 18 months.

- The suppliers are not authorized to supply material directly to any state Govt. / Semi Govt. / any other organization on the rate lower than the rate contract.

- The supplier shall supply the goods directly to the indentor / purchaser at the address given in the supply order.

- The supplier shall raise bills directly in the name of indenting officer / purchaser against the supplies made directly by them to the indentor’s satisfaction in compliance with the conditions contained in the supply order.

- The supplier shall receive payment against its bill directly from the indenting department / purchasing department. In case of Non-payment for the supplies made by supplier, they will demand payment directly from the department / indentor concerned and in no case Purchase Committee shall be responsible for the consequence for delayed payment or non-payment.

- The supplier shall carefully read all the conditions of tender for supply of equipment, floated by the Purchase Committee, and accept all terms and conditions in the tender document. Signing this contract means that the supplier has read all the terms and conditions and abide by it.

IN WITNESS whereof the parties hereto have caused this agreement to be executed in accordance with their respective laws the day and year first above written.

That, in token of this agreement, both parties have today affixed their signature at Indore.

Signed, Sealed and delivered by the

Said ……………………………… (For the RATE CONTRACTING AUTHORITY)

In the presence of: …………………………………..

Signed, Sealed and Delivered by the

Said ……………………………… (For the supplier)

In the presence of: …………………………………..

ANNEXURE-IX

Ref. Clause No. 13.3 (j) of ITB

DETAILS OF MANUFACTURING UNIT / AUTHORIZED DISTRIBUTORS

Name of the Tenderer & Full Address 	:

(Whether manufacturer / authorized distributor)

PAN number 				:

Phone Nos. 				:

Fax No.				:

E-mail Address				:

Date of Inception 				:

Equipments Manufacturing / Distribution License No & Date 	:

Issued by 				:

Valid upto 				:

CST / VAT Registration No.			:

If bidder is authorized distributor then 	:

name, address, telephone, fax of

authorized manufacturer.

Name & Designation of Authorized Signatory

Signature of the Authorized Signatory

The details of manufacturing unit / authorized distributor shall be for the premises where items quoted are actually manufactured / stoked.

ANNEXURE-X

Ref. Clause No. 17.2 of ITB

PRICE SCHEDULE

| Sr No | Item Code | Name of the Equipment / Item | Name of Manufacturer | Make & Model No. | Rate per Unit (Landed price) (Inclusive of excise / custom duty, transportation, insurance, service charges, inspection charges and any incidental charges etc.) | Amount of Transportation, Insurance, Service charges, Inspection charges (included in quoted rate per unit) | Rate of Excise / Custom Duty (included in quoted rate per unit) | Rate of CST/ VAT as applicable |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (1) | (2) | (3) | (4) | (5) | (6) | (7) | (8) | (9) |

Place:

Date:

Signature

Name in Capital Letters

Designation

Note: This format of price schedule is a sample for the Bidders. Price schedule should be submitted online only in the prescribed format and price schedule should not be submitted in physically or in form of CD, otherwise bid shall be rejected.

ANNEXURE-XI

Ref. Clause No. 17.2 (vii) of ITB

PRICE SCHEDULE FOR ANNUAL (WITHOUT SPARE PARTS) /COMPREHENSIVE (INCLUDE FREE LABOUR, REPAIR, OTHER SERVICES & SPARE PARTS) MAINTENANCE CONTRACT

(A.M.C. / C.M.C. ) AFTER EXPIRY OF WARRANTY

(RATES SHOULD BE QUOTED IN PERCENTAGE OF THE VALUE OF THE MACHINE)

| Sr No | SME Code No. | Name of the Equipment | For first year with spare parts & labour | For second year with spare parts & labour | For third year with spare parts & labour | For fourth year with spare parts & labour | For fifth year with spare parts & labour |
| --- | --- | --- | --- | --- | --- | --- | --- |
| (1) | (2) | (3) | (4) | (5) | (6) | (7) | (8) |

Place:

Date:

Signature

Name in Capital Letters

Designation

Note: This format of price schedule is a sample for the Bidders. Price schedule should be submitted online only in the prescribed format and price schedule should not be submitted in physically or in form of CD, otherwise bid shall be rejected.

ANNEXURE-XII

Ref. Clause No. 13.3(g) of ITB

DETAIL OF SERVICE CENTER IN M.P.

| S. No. | City | Name & Place of Service Center | Address, Telephone, Fax & email | No. of Service Engineer with Name / Mobile No. | Remark |
| --- | --- | --- | --- | --- | --- |
| 1 | Indore |  |  |  |  |
| 2 | Bhopal |  |  |  |  |
| 3 | Gwalior |  |  |  |  |
| 4 | Jablapur |  |  |  |  |
| 5 | Rewa |  |  |  |  |
| 6 | Sagar |  |  |  |  |

Name & designation of the authorized Signatory

Signature of the authorized signatory

CHECK LIST FOR TERMS AND CONDITIONS FOR EQUIPMENTS

Check list for Terms and Conditions  (To be filled by the bidder and submitted along with the bid) Page No. must be mentioned against each serial.

All non notarized documents except EMD must be self attested together with official seal .

| Sequence for online upload of document | Documents | Check Yes / No | If, Yes Page No. |
| --- | --- | --- | --- |
| Upload 1 | 1. Annexure – I (Sales Tax Clearance Certificate) |  |  |
| Upload 2 | 2. Annexure-II (Manufacture Authorization Form) |  |  |
| Upload 3 | 3. Annexure – III (Declaration / Undertaking Form) |  |  |
| Upload 4 | 4. Annexure – IV (Proforma for Performance Statement) |  |  |
| Upload 5 | 5. Annexure – V (Annual Turnover Statement) |  |  |
| Upload 6 | 6. Annexure – VI (Specification of require Equipments) with two additional self certified copies duly signed and stamped on each page |  |  |
| Upload 7 | 9. Annexure – IX (Details of Manufacturing Unit) |  |  |
| To be filled online only | 10. Annexure – X (The Price Schedule) |  |  |
| To be filled online only | 11. Annexure – XI (Price Schedule for AMC/CMC) |  |  |
| Upload 8 | 12. Annexure – XII (Details of Service Centre in M.P.) |  |  |
| Upload 9 | - Registration Certificate of the company with details of the Name, Address, Telephone Number, Fax Number, e-mail address of the firm and of the Managing Director / Partners / Proprietor. |  |  |
| Upload 10 | - Authorization letter from manufacturer authorizing a person to transact a business with R.C.A. |  |  |
| Upload 11 | - The instruments such as power of attorney, resolution of board etc., authorizing an officer/person of the bidder should be submitted with the tender and such Authorized officer/person of the bidder should sign the tender documents. |  |  |
| Upload 11 | - Market Standing Certificate issued by the Licensing Authority as a Manufacturer / distributor for each equipment quoted for the last 3 years. In case of direct importer, evidence for importing the said items for the last three years. |  |  |
| Upload 13 | - The bidder should also submit national & international quality certificates like ISI/CE/C ISO-9002, IP/BP etc” mark / IEC standard or equivalent certificate of quoted product, if available. |  |  |
| To be submitted physically only | - The bidders have to submit name of the items, its code no. for which they are quoting in the price bid. Such names and items code of the items should be submitted along with the technical bid, failing which the tenderer’s price bid will not be opened. The Bidder has to submit Name of Item and its code number. |  |  |
| Upload 14 | - Concern / Company has (a) not been debarred / blacklisted in the past either by Rate Contracting Authority or by any State Government or Central Government Organization. (b). firm has no vigilance enquiry / CBI enquiry pending against him / supplier. (c) The firm has not supplied the same item at the lower rate than quoted in the tender to any Govt. / Semi Govt. or any other organization. Affidavit to this effect shall be submitted by the concern / company and should be duly notarized on 100/- Non judicial stamp paper. |  |  |
| Upload 14 | - Original Bid Form duly signed by authorized signatory as per Section V, duly sealed and signed by the bidder on each page for acceptance of Terms and Conditions. |  |  |
| Upload 15 | - Bidders should have the registration under Commercial Tax Authority Registration should be attached. |  |  |
| Upload 16 | - An affidavit of the manufacturer on a Non judicial stamp paper on Rs. 100/- binding itself for supply of genuine spare parts to the purchaser for a period of minimum 5 years beyond warranty period. |  |  |
| Upload 17 | - An affidavit of the manufacturer on a Non judicial stamp paper on Rs. 100/- that the quoted equipment is the latest model of the company matching to the specification of this tender and that it is not refurbished equipment. |  |  |
| Upload 18 | - Certificate for being in business or more than 3 years |  |  |
| Upload 19 | - Certificate for sole ownership / partnership. |  |  |
| Upload 20 | - Statement of good financial standing from bankers. |  |  |
| To be submitted physically only | - The printed original catalogues of primary manufacturer and any other technical documents like data sheet or operational manual of equipment with highlighting the features in portal along with the other documents. In catalogue, the quoted product no. and name should be highlighted, against which that product is quoted. These documents are also to be submitted in physical form before due date along with Bid security. |  |  |
| Upload 21 | - In case of imported equipment IEC certificate of importer / bidder shall be submitted. |  |  |
| Upload 22 | - A separate price list of all spares and accessories (including minor) required for maintenance and repairs in future after guarantee / warrantee period. |  |  |
| To be submitted physically only | - Recurring expenditure on equipments. |  |  |
| Upload 23 | - Any other document required as per tender document. |  |  |

ITEM CODE MOT 1 : ESTABLISHMENT OF MODULAR OPERATION THEATRE

TECHNICAL SPECIFICATION -

The  main  requirements  for  any  Operation  Theatre  wall  system  are  that  they  should  be  impact resistant,  show  resistance  to  solvents  &  chemicals,  prevent  colonizing  of   micro-organisms  or neutralize micro-organisms which come into contact with the wall surface.  The surfaces should be stable – prevent cracking and movement, scrub able – amenable to  cleaning & have a completely sealed  finish,  Biological  attack  resistance  &  have  hygiene   finish  &  lastly  have  Hydrothermal performance.

•	All wall mounted equipment should be flush-mounted and sealed into theatre wall by means of durable sterile jointing system.

•	The  wall  panel  design  and  construction  should  allow  for  the  installation  and  support  of  all equipment’s with  provision  of openings  required for repair and  maintenance  without  affecting rigidity and strength. In case require, one should be able to easily remove the wall panel for service requirements.

•	The total distance between the inside and outside surfaces of the theatre should vary to suit the architect’s layout; however the composite wall thicknesses should be 100mm or less.

•	The cavity between the inner and outer walls should be left with minimum obstructions for the possible addition of equipment at a later date and to enable services, pipes, conduits etc, to be run within the cavity.

•	The Sterile coating is water resistant, does not support bacteriological or fungicidal growth and is resistant to most chemicals commonly used in hospital departments. It should neutralise micro-organisms which come into contact and should prevent colonizing or sporing of micro-organisms.

•	The  sterile  coating  shows  resistance  to  a  wide  range  of  mixed  species,  including  stubborn pathogens   such   as   MRSA,   Acinetobacter   sp,   Aerobacter   aerogenes,   Bacillus   subtitles, Escherichia  coli,  Listeria  monocytogenes,  Pseudomonas  aeruginosa,   Pseudomonas  putida, Salmonella typimurium, Serratia marcescens, Staphylococcus Aureus.

•	The sterile coating should remain unaffected by radiation, its anti-microbial system should not leach out & its anti-microbial properties should last for a minimum of 10 years.

•	The sterile coating system should be easily maintained and should withstand repeated cleaning with alkaline detergents, antiseptics and fumigation agents without any degradation to the surface finish or performance.

•	Conform to standards HTM 08-01 / HTM 56.

- PRE FABRICATED MODULAR OPERATION THEATRE - WALL, CEILING & SLOPING PANELLING

1.1  WALL  &    CEILING  PANELLING:

(All civil work related to Modular OT is in the scope of the Contractor)

The Pre-fabricated operating room should be free standing structure, constructed from composite, free standing insulated steel wall panels.

These panels should be produced in a double band laminator, in which two continuously moving belts of Stainless Steel 304 sheets of 0.80mm thickness are firmly bonded together by the  sprayed in insulating foam, which hardens during this process. All panels should have a fifteen year de-lamination warranty.

The core should consist of rigid polyurethane foam, which has been injected under high pressure, with a minimum density of 40 kg / m3. The walls should be modular in construction, consisting of 1180 mm, 875 mm, 570 mm and 265 mm wide panels, together with corner panels 200 x 200 mm to achieve the desired false ceiling height. The thickness of the complete panel, including the core should be a minimum of 60 mm. 	Conform with respective certificate copy duly attached herewith having clearly mentioning of above.

The modular wall panels are installed upon floor using a “U” channel. Each Pre-Fabricated Operating Room may optionally include sloping ceiling panels at an angle of 45 degrees from the straight wall panel edge. These options reduces the overall volume of the room, decreasing the running costs for air-conditioning and ventilation, while at the same time maintaining the necessary height at the center of the room for the installation of ceiling mounted equipment’s.

The inner surface of the operating theatre should be seamless, free from visible joints and sharp edges. All internal corners and panel joints should be filled with proper Epoxy filler, Sanded flush or ready to use on site, ready to receive the sterile plastic finish. All panel joints should be invisible when the final plastic coating is applied.

All wall-mounted equipment should be flush mounted and sealed into theatre wall by means of a sterile jointing system. The wall panel design and construction should allow for the installation and support of all equipment and the provision for openings required for the installations, without affecting the rigidity and strength.

The internal surfaces of the theatre walls and ceiling should be sprayed with water based, non-reflective liquid plastic, to a color approved by the architect, to a minimum dry film thickness of 60 microns. The plastic coating should overlap the floor covering, ceiling system and door frames by 25 mm to provide a continuous sealed surface.

The sterile coating applied to a thickness of 60 microns of the already applied primer coating on the panels inner surface, should be water resistant, does not support bacteriological or fungicidal growth and should be resistant to most chemicals commonly used in hospital departments.

The sterile coating should remain unaffected by radiation and other ionizing radiation at levels in excess of 1,000 mrad and should be classified to class 1 when tested in accordance with the requirements as specified under BS:476:part 7 ;1971, surface spread of flame test for materials. The coating should be SILVER Ion based.

The coating system should be easily maintainable and should be able to withstand repeated cleaning with alkaline detergents, antiseptics and fumigation agents without any degradation to the surface finish or performance.

The thickness of the wall panels should be a minimum of 10mm and mounted on the existing Wall structure using the necessary steel frame work. The Wall panels should be easily removable for catering to service related needs.

1.2 Electrical Components

Each modular panel should contain separate cable pathways for the routing of electrical services and a variety of openings and rear enclosures for the fixing of electrical components. Cable sizes for power sockets, earth and potential equalization wiring should be provided according to the specific requirements of the site. Separation of LV & ELV Supplies to BS7971, conforming to HTM 06-01 & HTM 06-02 standards.

1.3 Communication Components

Each  modular  panel  should  contain  separate  cable  pathways  for  the  routing  of  communication services and a variety of openings and rear enclosures for the fixing of components. Internal wiring is executed according to the existing valid standards, running in a separated cable pathway up to an integrated services connection point.

1.4 Mechanical Components

Medical gas outlets are connected to the wall panels from pipes running within. Piping for medical gases is segregated by means of a separate duct within the modular panel; the piping used will be medical grade copper tube according to EN737-3. The modular panel is delivered in a pre-piped configuration, including the CE-marked medical gas outlets of which all common brand types are available.

- ESD VINYL FLOORING

The flooring should conform to standards DIN EN 1081, DIN EN 1815, EN 12466, EN 425, EN 423 & EN 433. It should be fire resistant, scratch resistant, chemical resistant & resistance to fungi – bacterial growth.

Vinyl is the oldest effective ESD flooring material. Solid vinyl tile (SVT) can give a facility that hospital look of cleanliness and shine. Facilities managers often favor vinyl floor because vinyl has a reputation for being inexpensive, simple to repair and easy to clean. The joints between individual tiles are to be welded by heat fusion process to get a seamless floor. The PVC welding bar should of the same matching color as that of the tile and a hot air gun is used for fusion of welding bar with flooring tile. Conductive flooring to be installed within the OT area to prevent damage to sensitive medical electronic equipment’s, inadvertent movement of surgeons / technicians, accumulation of static dirt. Maintenance should be simple such as each individual tile can be replaced in very short down-time. The conductive flooring should ensure:

•	The protection of Electronic equipment and components

•	The protection of people against electro-static discharge risks

•	The protection of premises against explosive risks.

The Conductive tiles should be static – conductive, flexible, homogenous vinyl floor covering with an electrical resistance of 5 x 104  ≤ Rt ≤ 106  (EN 1081).

The Conductive flooring should have the following characteristics:

•	Low accumulation of electrostatic charges to people & equipment.

•	Should provide an enduring resistance to static and dynamic loads.

•	Should display excellent resistance to chemical products such as detergents, acids and alkaline products.

•	Low VOC emissions.

•	Should have Fungi-static and bactreo-static treatment throughout the total thickness of flooring.

•	It should be non-absorbent, impervious and non-porus.

- IMPORTED PLAN-AIR VENTILATION CEILING / PLENUM / AIR & LIGHT DIFFUSER / LAMINAR FLOW – CONFORMING TO STANDARD HTM 03-01 & HTM 60

The Plan-air ceiling ventilation system should dilute the bacteria generated by the operating team and patient in the theatre and to create an Air flow pattern that carries contaminated air away from the operating table & entering surgical wounds. It should be designed such that filtered, sterile air flows through the operating zone without an admixture of indoor air. The germs and aerosols released by occupants are displaced into adjacent room zone and removed with the exhaust air.

The main components that the Laminar flow should be having is a rectangular air outlet housing with air outlet frame and air discharge element on the underside and two housings on the top side each with a built in HEPA filter. The air discharge element should be a fine - meshed fabric for laminar displacement flow. The air discharge element should be split with a feed through for the surgical light. The element should be removable or can be folded downwards for easy accessibility to the housing interior for cleaning and disinfection. It should be also easily accessible to the HEPA filters which can be easily replaced when required to do so.

The complete Laminar flow system should be Pre-manufactured and pre-assembled unit, modular in design, having connection box and filter frame. The construction should consist of four units’ modular in design with filter frames and variable connection hoods for the supply air, the air discharge element should be made of fine-mesh laminar fabric (polyester) with a surrounding stainless steel frame. The air outlet housing should be made of stainless steel and the filter housing of powder coated steel sheet. The Laminar air distribution element should be modular in design, divided into four parts with a minimum number of bars, individually foldable, covered with monofilament polyamide from one side, fixed without screws. It should have a connection for differential pressure measurement.  The Laminar flow should be of following characteristics:

- Plan-air Pleanum Size : 2800mm x 2800mm with 4000CFM capacity Laminar Air Flow

- Plan-air Pleanum Size : 1800mm x 1800mm with 2000CFM capacity Laminar Air Flow

The system should be certified according to DIN 1946-4 (12-2008), tested to DIN 4799, having horizontally arranged terminal HEPA filters of H 14 type with grip protection on both sides; filter exchange from pure air side, filter tested and certified according to DIN EN 1822-1.

- DOOR AND FRAME

Hermetically Sealed Sliding Door – Automated

- Automated operated hermetically sealing sliding main door on OT, of appropriate size,

2100mm (H) X 1000 (W) and 2100mm (H) X 1500 (W).

- Having a track system wherein the door blade is guided under a 45° angle with approximately 6mm deep indentations at the closed position.

- Having a special inside & outside Stainless Steel lever handle that lifts out the door blade from its hermetic  closing  position,  thus  allowing  even  very  heavy  doors  to  be  opened  effortlessly, manually.

- Encased within an extruded anodized 2.5 mm thick aluminum profile frame all around the door blade with solid high pressure laminate on both sides.

- Should have special nylon top rollers with double roller bearings, to enable the door to slide smoothly & silently.

- The door blade to have a special 3-tier specially designed 3 side heavy duty replaceable EPDM gasket, for the hermetic sealing on the wall frame. The bottom sealing should be with a 2-tier heavy duty EPDM gasket to flush with the finished floor.

- The door core should be 60mm thick, built up with 4mm thick HPL (High pressure compact laminate) skin on both sides of the door to provide better strength and rigidity. Having between the two sides of laminates – CFC free, high density polyurethane puff (density 40 Kg / m3), thickness 52mm.

- Vision panels - double glazed, laminated safety glass, flush finished on both sides with door blade should be provided in the doors.

- The automation controller for the door should be microprocessor based with regulated electro-mechanical sliding door drive.

- The automation should be able to operate a door with maximum door weight of 250Kg.

- The wall frame should be high grade aluminum extrusion with natural anodized 15 micron finish wall frame profile. The wall frame profile should be 3 sided blind fixed with cladding on wall cut-out section from inside and outside of the room.

- The automation should have variable / adjustable opening & closing speeds, with opening speed adjustable between 120mm/sec --- 800mm/sec and the closing speed  adjustable  between 120mm/sec ---- 500mm/sec.

- The automation / door can be controlled/operated using either push-button switch, Kick switch or Sensor (touch less – infra-red) switches OR all in combination.

- Double Leaf SS Door

Double Leaf Stainless Steel Door Single Side Open Flush Door with 49mm thickness of both door leafs, having size of 1800mm x 2100mm with view window of 300mm x 300mm duly complete with clean view glass, The Door should be having 0.8mm thick SS304 sheet on both faces while door core should be of PU, the Door should be hinge mounted and should complete with all necessary accessories i.e. Door closer, D type SS Door Handle etc.

- Single Leaf Stainless Steel Single Side Open Flush Door

Single Leaf Stainless Steel Single Side Open Flush Door with 49mm thickness, having size of 1000mm x 2100mm with view window of 300mm x 300mm duly complete with clean view glass, the Door should be having 0.8mm thick SS304 sheet on both faces while door core should be of PU, the Door should be hinge mounted and should complete with all necessary accessories i.e. Door closer, D type SS Door Handle etc.

- WALL MOUNTED ITEMS:-‐

Operation theatre control panel

It should be a digital control panel consisting of a computer and a touch screen managed by a PLC. It should integrate all controls, signals and alarms of the operation theatre and offer several services addressed to the maintenance staff.

The information management system should be managed by the programmable software – PLC and the visual and control parameters by the PC / Touch screen, IP64 protected, 15” (inches) in size. It should have a digital microphone and speakers, a 300K pixel camera, ambient light sensor and a wireless LAN 802.11 a Bluetooth V2.0.

Following frequent used functions should be available:

a)   Controls:

	General Light Switch control, 3 controls in number, having dimming facility

	Temperature and humidity display

	Operation theater status – occupied / free / cleaning

	Stopwatch control

	OT surgical light switch (depending upon the make, model of the light)

b)   Alarms:

	Electrical Alarm (visual & Acoustic)

	Gases Alarm repeater (Visual & Acoustic)

	UPS charging status

c)   Communications:

      Telephone / IP Telephone if line available

      Ethernet connection

      SMS Alarms

      Video – Conference facility

d)   Signals:

	Digital Inputs / Outputs: 14 Nos / 10 Nos.

	Analogue: (0-10V) Inputs / Outputs (2 Nos. / 1

- X-ray viewing screens (two plates)

- The theatre should be equipped with a twin plate X-Ray Viewing Screen, recessed in the wall and designed to provide a high level of control luminance, without flicker, from a unit that is easy to clean and maintain.

- The  X-Ray  viewing  screen  illumination  should  be  by  LED  lamps,  controlled  by  dimming ballast.

- The  front  panel  diffuser  is  of  a  glare  free  type,  sealed  flush  with  the  inside  face  of  the operating theatre wall (or may as an option be integrated within the control panel fascia).

- The LED lamps should provide a uniform level of illumination across the entire front panel.

- Access for maintenance and lamp changing is from the front of the panel.

- All internal wiring is terminated in connectors with screw and clamp spring connections.

- Individual fuses or miniature circuit breakers protect all internal circuits.

- All internal wiring is of high temperature resistance and secured by propriety cable clips.

- It is flush mounted and sealed into theatre wall by means of a sterile jointing system.

- Equipment Storage Unit

- Equipment Storage Unit to be provided in operating theatre of dimensions approved by the user.

- It should be flush mounted into the theatre wall with a sterile jointing system.

- The cabinets should be made of Stainless Steel AISI 304 material.

- Doors of cabinet should open right to left and made of clear safety glass.

- It is suitably partitioned / shelved to allow storage of common endoscopic       equipment’s.

- It has the facility of secure locking.

- Conform to standard HTM 63

- Operating List - Writing Board

The Operating List – Writing Board should be flush mounted with a sterile jointing system, The Color of Writing Board should be White on front surface for dry wipe marker and green back for chalk, used material should be Rubber Magnet + Paper, the Operating List-Writing Board can bear Temperature upto 800◦ Celsius, the used sheet should be of Triple Acrylic Resin Coated Imported Steel Surface, while the writing board should have surface being 0.40mm.  It attracts magnet, Smooth & One Stroke Erasing surface, Writing Board surface should have Chemical treatment surface totally Rust free/Crack Proof/Stain Resistance and should have Eco Friendly characteristics.

- Cascade Pressure Stabilizer

•	Cascade pressure stabilizers should be a range of multi-bladed units specifically designed  to control room air pressures in critical areas, such as operating theatre suites.

•	Each stabilizer should comprise of carbon steel housing with up to four Grade 304 Stainless steel blades, which pivot upon sealed for life bearing assemblies.

•	Balancing is carried out utilizing a proven balance weight assembly.

•	These stabilizers should be accurate to a range of 1 Pascal over their working range of 5-35Pa.

- Hatch Box

A hatch should be provided to remove waste material from the operation theatre to dirty linen area just adjacent to the operation theatre. Each hatch should be equipped with two doors and the door should be operated electronically. The hatch should be designed in such a way that only one door should be opened at one time. The UV light should be so installed that it is kept on while both the doors are closed, this UV light should automatically turn “OFF” in case of opening of either of the doors. There shall be indicators on both sides of the OT such that door open / close status can be monitored from both ends.

- Clean-room Illuminators

Lamp covered with translucent Plexiglas with high transmission level (higher than 80%), held  in  a frame of white shell sheet, with integrated screw angles for fastening the lights to the panel system. Including closed mounting frame for the built-in LED lamps; the connection between the illuminator and the ceiling hollow space should be completely airtight. Clean-installed light should non-hygroscopic and be of IP 54 grade. Dimensions should approximately be 600 x 300mm (length x width).

- Surgical Scrub Sink – Two Bay

The Surgical Scrub Sink should be made of Stainless Steel of grade AISI 304. It should be destined to be used for operating theatres, where high hygienic standards are to be maintained. It should be free standing sink with wall mounted possibility, having a depth of 300 mm. It should be two bay, having standard  2  x  drain  traps,  water  and  soap  dispensers  activated  by  sensor,  water  and  soap  time adjustable, water thermostat – temperature adjustable and having chrome  faucets. It should have open able front panel with two latches for maintenance purposes, having approximate dimensions of 1600 (L) x 635 (w) x 1271 (H) in mm.

- HVAC System

15.1 AHU - AIR HANDLING UNIT (DOUBLE SKIN TYPE)

Type

The air-handling units are of double skin construction, draw-thru type comprising of various sections such as Pre-filter section, coil section, as shown on drawings and included in schedule of quantities. Unit must be able to work satisfactory in exposed atmosphere conditions.

Casing

Double skinned panels are fabricated with anodized extruded aluminum extrusion frame work bolted together with sandwich panel having powder coated 0.70 mm sheet for outer skin and plain GP 0.63 mm sheet for inner skin. 43 mm thick PUF insulation material is injected between the two panels (with U value no greater than 0. 85 W / m2/k)

The entire frame duly painted is mounted on sheet steel channel based.  The panels are sealed to the framework  by  heavy-duty  'O'  ring  gaskets  held  captive  in  the  framed  extrusion. All panels are detachable or hinged.  Hinges are made of die cast aluminum with stainless steel pivots, handles are made of hard nylon and be operational from both inside and outside of the unit.  All fixing and gaskets shall be concealed.

Motor and drive

Fan motors are highly efficient and work on 440 ± 10% volts, 50 cycles, three phase with explosion proof type with class F installation, with IP 55 protection.   Motors are especially designed for quiet operation and motor speed does not exceed 1440 rpm.   Drive to fan is provided through belt-drive arrangement.  Belts are of the oil-resistant type.

Fan

Fan are of centrifugal type, conforming to AMCA 210 and are double width, double inlet with forward- inclined airfoil blades, specially designed and suitable for the required operating pressure.  Fan casing are made from galvanized steel sheet.   Fan shaft is grounded C 40 carbon steel and supported in self-aligning plumber block operating less than 75% of first critical speed, grease lubricated bearings.

Cooling Coils

DX coils have 12.5 to 15 mm dia tubes minimum 24G thick with sine wave aluminum fins  firmly bonded to copper tubes assembled in zinc coated steel frame.  Face and surfaces areas are such as to ensure rated capacity from each unit and such that the air velocity across the coil does not exceed 150 meters per minutes.  Each coil is factory tested at 21-kg/m2   air pressures under water.   Tube is Hydraulically / mechanically expanded for minimum thermal contact resistance with fins.  Fin spacing is 4 - 5 fins per cm.

Filters

Each unit is provided with a factory assembled filter section containing washable synthetic type air filters having anodized aluminum frame.   The media is supported with HDP mesh.   Filter banks are easily accessible and designed for easy withdrawal and renewal of filter cells.   Filter banks face velocities do not exceed 100 m/minutes. Differential pressure switch is to be fixed across the filter as part of AHU's system.

Ducting

The duct supply system should be free of construction debris. Ducting shall be made of Aluminum with curves & bends where indicated for easy flow of air and ensured to be air tight by applying silicon sealant after fabrication. Hangers shall be provided to ducts & shall be suspended by means of G.I. coated rods & these shall not be more than 2.5mtrs apart.  Thermal insulation with 9mm XPE for supply & return air ducts. Joints will be lapped with Nitrile rubber tape for better insulation.

15.2 Air Cooled Package Chiller / Condensing Units.

Compressor

The compressor shall be of screw, scroll / reciprocating type, hermitic, in accordance with ARI 520, direct driven with capacity control arrangement. The compressor casing shall be of cast  iron and designed for 450 psig or higher.  The compressor shall incorporate rolling element bearing to support rotating assembly. The rotor shall be higher steel alloy.  We prefer the use of Imported Compressors Hermetic Compressors.

Refrigerant  circuit  components  shall  include  flexible  pipe  connectors,  hot  gas  muffler,  high  side pressure switch, liquid line shut-off valves, suction and discharge shut-off valves, filter drier, moisture- indicating sight glass, electronic or thermostatic expansion valve (EXV), heavy duty pressure gauge with cocks to monitor suction, discharge and oil pressure and complete operating charge of refrigerant and compressor oil.

Motor

Compressor motor shall be of Hermetic / semi-hermetic direct drive, squirrel cage, two pole, induction type, refrigerant cooled motor suitable for 415V / 50 Hz. 3 phase supply.  Hot gas motor cooling is not acceptable. We prefer the use of Imported Compressors Hermetic Compressors.

Condenser

Condenser shall be Air Cooled type. Tubing shall be copper, Aluminum fins high efficiency type.   Tubes shall be nominal 19mm.  Outer diameter and thickness shall not be less than 22 g.  and  rolled  into  rube  sheets  and  shall  be individually replaceable and also tubes shall be coated with corrosion resistant coating.

Condenser fans shall be direct coupled to motor and protected against overloading and with minimum 1.15 service factor.

Condenser shall design to allow isolating refrigerant charge when servicing the compressor.

With maintain the above said technical specifications, The Air Handling Unit with Air Cooled Package Chiller / Condensing Units should have following capacity:

- 17.5 tr with 5000CFM

- 11 tr with 3500 CFM

- Electro Hydraulic OT Table

C-Arm Compatible Insulated Operation Table made with whole table top made up of radiolucent material and automated table top movement, is able to bears a maximum of 200 Kg patient weight with hydraulically driven feet. The Operating Table is having power supply of 24V SELV (Safety Extra-Low Voltage) & complies completely with safety requirements of IEC 60601-1. The OT Table top is constructed modularly, into 4 distinct segments – leg plates, seat plate, back plate and head plate to suite various surgical disciplines with minimum table top height can be as low as 600mm.Table top padding provides good air circulation and prevents moisture ingression with double-layered table top mattress, an elastic layer uniformly distributes patient weight over the largest possible surface and protects patients from the hard table top, U shaped base of the Operation table provides extra room for surgeons to stand in case of perineal operations, beside the usual automated movements, back plate movements and longitudinal shift are also automated and desired position is achieved by simply one touch. The OT Table comprises of following characteristics, Table Size (Full length -  2060mm x Height without padding is 600mm – 950mm x Width without side rail is 520mm x Width with side rail is 590mm.

| S.No. | Technical Specifications |
| --- | --- |
| 1 | Table Top Length : 2100mm |
| 2 | Table Top width : 500mm |
| 3 | Table Height Adjustment Range: 700mm-1030mm |
| 4 | Longitudinal Shift : 296mm |
| 5 | Trednelenburg / Reverse Trednelenburg : ±25° |
| 6 | Tilt Left / Right : ±15° |
| 7 | Back Rest Adjustment: +65° / -40° |
| 8 | Head Rest Adjustment : +30° / -60° |
| 9 | Leg Rest Adjustment : Down 90° |
| 10 | Load Capacity : 135kg |
| 11 | Power Supply : AC 220v / 50Hz, AC 110V / 60Hz |
| 12 | Input Power : 200VA |

- All Wiring within the Modular Operation Theatre

Wiring with Low Leakage Current wires of FRLS wires including providing and fixing of PVC conduit and boxes etc. to complete the work in all respects. Wiring for 250 volts single phase & neutral 6/16 amps switches socket with 4 Sq mm PVC insulated copper conductor 1100 volts grade standard flexible wires in concealed surface mounted 25mm dia 16 gauge PVC conduit including providing and fixing of 6/16 amps 5 pin switch socket of approved make and design in 1mm thk. G.I. box with grid plates and earthing the third pin of each socket outlet with 4 Sq mm PVC insulated copper conductor trended flexible wire. Provision for UPS wiring and sockets within the OT.

- Medical Gas Pipeline System within Operation Theatre

Medical Gas Pipe Line System inside the Modular Operation Theatre for two pendants only completed in all manners.

- View Window

View Window should be flush mounted in Operation Theatre wall, and should have the Size of 1000mm x 1000mm. The Window should comprises Remote controlled Motorized venetian blinds of powder coated aluminum strips of vista level or equivalent, for rigidity, the Glass of the window should be laser cut toughened  Glass of 22mm.  The window should be sealed with rubber gasket.

20    OT LED light

- The double dome operating light must be designed for the use in high demanding surgical procedures. State-of-the-art LED bulbs should be used to ensure a low energy consumption and a long service life.

- Revolving outer handles at the light head should be provided to allow for non-sterile positioning.

- Light head must be designed with smooth transitions and surfaces, without slots, gaps or exposed screwing to ensure fast and effective cleaning.

- The light head with streamlined shape is favourable within laminar flow. The light head must be resistant to disinfectant.

- For sterile positioning an ergonomic, exchangeable and centrally positioned sterile handle within the light head should be provided.

- For visualization of the surgical procedure the light system should be ready to mount with an HD video camera system in future.

- All main joints of surgical light must be provided with unlimited rotation (360°). Light head and suspension must be sealed dustproof.

- All domes should have Illumination intensity of endo light - 3,000 lux

- Central axis should have facility to mount 3 domes and one arm for monitor AC 3000 with swivel arm. VESA mount.

- Bulbs should lose only 20% of their light output after 30,000 hours

- Classification

  - Protection class acc. to IEC 60601-1  -

    - Multipurpose switching power supply - Protection class 1

    - Light body – SELV

  - Arm system – IP41

  - Light body – IP42

  - Ventilation and air conditioning Part 4 - DIN 1946-4:2008-12

  - Electrical installation - IEC60364-7-710/VDE0100-710

  - Product design - IEC60601-2-41/IEC60601-1

  - Electromagnetic compatibility (EMC) - IEC 60601-1-2

  - Classification as per EC Directive 93/42/EECAnhang IX - Class I

  - UMDNS-Code (Universal Medical Device Nomenclature System ) - 12-347

- The surgical light should be complete with all components for ceiling mount and electrical feed-in, incl. finalised installation.

- Technical data for main dome: - 1nos.

    - Central illumination intensity Ec			 160.000 lux

    - Light field diameter adjustable                                     220 mm to 290 mm

    - Focusing                                                                             mechanically

    - Depth of illumination L1+L2 				1100 mm

    - Active light emitting surface				4893 cm2

    - Average Color rendering index Ra			93

    - Color rendering index R9 (red)			94

    - Color temperature more than                                        4500K

    - Central illumination at 1m distance with:

      - Tube						99%

      - one mask:					70%

      - tube and one mask:				70%

      - two masks:					50%

      - tube and two masks:				50%

    - Total illumination intensity of a double light combination 	<1000 W/m²

    - Energy efficiency Ee - 					540 W/m²

    - Ee/Ec ratio - 						3.4 m W/m² x lux

    - Photometric radiation equivalent - 			280 lm/W

    - Adjusting the illumination intensity			36 % to 100 %

    - Light-emitting surface					3866 cm²

    - Bulb efficiency						60 lm/W

    - Total nominal power of all LED bulbs			108 W

    - Maximum power consumption per light			200VA

    - Supply voltage					100-240V 50/60 Hz

    - Light head supply voltage					24V/DC

    - Rotation Radius					360°

    - Lifetime of light sources				30000 h

- Technical data for satellite dome -1nos. (dome should be capable to mount camera in centre of light head)

    - Central illumination intensity Ec			110.000 lux

    - Light field diameter adjustable                                     220 mm to 290 mm

    - Focusing                                                                         mechanically

    - Depth of illumination L1+L2 				1300 mm

    - Active light emitting surface				2373 cm2

    - Average Color rendering index Ra			93

    - Color rendering index R9 (red)			94

    - Color temperature more than                                        4500K

    - Central illumination at 1m distance with:

      - Tube						99%

      - one mask:					50%

      - tube and one mask:				50%

      - two masks:					45%

      - tube and two masks:				45%

    - Total illumination intensity of a double light combination 	<1000 W/m²

    - Energy efficiency Ee - 				330 W/m²

    - Ee/Ec ratio - 						3 m W/m² x lux

    - Photometric radiation equivalent - 			280 lm/W

    - Adjusting the illumination intensity			42 % to 100 %

    - Light-emitting surface					2373 cm²

    - Bulb efficiency					60 lm/W

    - Total nominal power of all LED bulbs		60W

    - Maximum power consumption per light		150VA

    - Supply voltage				100-240V 50/60 Hz

    - Light head supply voltage				24V/DC

    - Rotation Radius					360°

Lifetime of light sources				30.000 h

- Technical data for HD Camera

  - Material for camera casing				aluminum

  - Frame chip						1/3“ CMOS sensor

  - Image resolution			full HD, 1920 x 1080 pixel

  - Refresh rate			PAL / NTSC / 1080i/50 / 1080i/60

  - Image signal output to control unit  2 x HD-SDI / 2xDVI/1xY/U/V (via 3y              				BNC)/1xFBAS (via BNC)

  - Video output of control unit

    - HD-SDI (2x)		1080i/50Hz / 1080i/60 Hz BNC jack

    - DVI (2x)			1080i/50Hz / 1080i/60 Hz DVI-D jack

    - CVBS			PAL/NTSC BNC jack

    - Y/C				PAL/NTSC mini DIN 4-pole jack

    - YUV				1080i/50Hz / 1080i/60 Hz BNC jack

  - Zoom			120-fold (10-fold optical / 12-fold									digital)

  - Lens length			3.4 mm / Tele 33.9 mm

  - Signal-to-noise-ratio	50 dB

  - Auto shutter			1/2 to 1/10.000 seconds

  - Minimal illumination level	12.0 lux

  - White balance		automatic / manual

  - Motorized image erector	> 360°, electromotive

  - Still frame / freeze		digital

  - Focus			automatic / manual

  - One-touch-autofocus	required

  - Iris (manual)			required

  - Auto-iris			required

  - Intensifier			automatic / manual (-3 to 28 dB, 16

levels)

21 - Ceiling Supply Unit Anaesthesia  Pendant

# 1.1 General Standard of Execution, Regulations

The corresponding requirements of the European Norm ISO EN 11197 (special requirements for the safety of medical supply units) Medical Device Directive (MDD) 93/42 EEC as well as the relevant rules and regulations of the law for technical work equipment (equipment safety law) and ISO/IEC (see following table) in their valid version must be fulfilled through the following specified equipment.

Qualified vendor must have a Quality Management System acc. to ISO 9001 and ISO 13485 (Medical) in place, as well as a MDD (Medical Device Directive) annex 2 section 3 certification

- 1.2 Individual Standards and Regulations

| Standards | Basic requirement of the Medical Device Directive (MDD 93/42/EWG) • DIN EN 60601-1 / medical electrical devices • DIN EN 60601-1-2 / medical electrical devices • DIN EN 793 / medical supply units • EN ISO 14071 / medical products – Risk analysis |
| --- | --- |
| Classification | Class II b acc. MD Directive 93/42/EWG appendix IX UMDNS Code: 18-046 |

The equipment must be CE-certified or has got the approval of the European regulatory authority according to MDD.

The quotation must include all facilities and equipment components, including necessary accessories which are required for the functionality of the equipment.

Variations or alternatives to the above-mentioned standards or to the following specified scope of performance must be specifically marked with the abbreviations ”Var.” or ”Alt”.

2.	Scope of Medical Supply Unit Equipment

The supply unit must meet the high standards for a safe and reliable operation in the OR.

The supply unit must have upgrading capability of new functions for adaptation to future requirements.

- All surfaces resistant against corrosion and disinfectants.

The functional units of the supply unit must be completely assembled and tested by the manufacturer.

## General Requirements for Ceiling Supply Units

Ergonomic, optimized supply unit consisting of the following functional units:

- Ceiling fixture set for installation to concrete ceiling structure.

- Ceiling interface for connection of supply unit with electrics and gases delivered from site.

- Horizontal swiveling arm system.

- Media column with terminal units for medical gas supply/gas evacuation,
high- and low voltage power supply according to specification.

- Workstation components such as shelves, drawers, extension arms and equipment holders for ergonomic workplace design according to customer requirements.

- Friction brakes for fixation of adjusted arm position.

- Complete unit including intermediate ceiling construction and installation.

The previously mentioned functional units must be coordinated mechanically, functionally and ergonomically, thus presenting a complete medical supply unit.

# Scope of Delivery- and Technical Performance

## Basic Units

### 3.1.1	CSU Light Load Arm System;
Arm Length 1000+500 mm
with Media Column  1000 mm

- Arm system with two horizontal swivel arms, intermediate bearing and ceiling bearing.
Action radius 1500 mm

- Load capacity on media column min. 120 kg

- Height of media column approx. 1000mm

3.1.2 Media Column

- Height of media column 1000 mm

- Fitting of installation components possible on 3 sides.

- Pivot bearing integrated into media column.

- Swivel range 320 - 330 °

- Adaptation to ceiling height via distance tube attached to arm system.

- Mounting system for attachment of workstation components, 	
e.g. shelves, drawers, side arms, equipment holders
from front side.

- Working height of workstation components infinitely variable on media column.

- Set of control elements, consisting of 2 handles for comfortable operation of the supply unit.

- Ergonomically designed handles to support user for positioning the system.

- One-piece front plates without gaps.

- Surfaces resistant against disinfectants.

## Installation Components

### 3.2.1	Gas Fittings

Media head / media column equipped with:

2_ x  O2	(oxygen)

1_ x Air 4	(compressed air 4 bar)

2_ x VAC	(vacuum)

1_ x N20 (Nitrous Oxide)

### 3.2.2	High voltage electrical installation

Media column equipped with:
8_ x electrical  6/16 Amp universal socket.

### 3.2.3	Low voltage electrical installation

Media column equipped with:
1_ x RJ 45 Socket.	
1_ x BNC socket.

## Workstation Components

### 3.3.1	Shelf for mounting at front rail system of Media Column

Shelf for attachment to front rail system of previously described media column - 3 nos.
Surface area 440 x 480 mm

(W x D)

- Load capacity min. 40 kg	
Thickness max. 22 mm

- Material: Aluminium

- Smooth, scratchproof, fade-resistant and antistatic surface.

- No sharp edges and grooves.

- Integrated rubber elements at all 4 edges for injury prevention.

- With curved, ergonomic design.

- Sida rails should be provided with each shelf.

### 3.3.2	Drawer for mounting at shelf – 1 Pcs.

Drawer for attachment to shelf, as previously described
Dimensions 440 x 340 x 170 mm (W x D x H)

- Load capacity min. 5 kg

Special drawer runner system with

- exceptionally smooth sliding function

- self-closing function

Facilitated cleaning by

- easy removal capability of complete drawer

- removable front fascia

- smooth surfaces without protruding handles

- Future exchange of front fascia possible due to retro-fit capability

### 3.3.3	CVP-Pole – 1 Pcs.

- Support tube with height adjustable extension pole

- Length of support tube min. 1000 mm

- Length of extension pole min. 800 mm

- 4 hooks, cross orientation, for accommodation of infusion bottles

### 3.3.4    Extension arm for mounting at front rail system of Media column

- for attachment of a small equipment pole Ø25 mm or CVP-pole

- with locking screw on each swivel joint

- unrestricted horizontal movement around support rail axis

- Length min. 300 mm

Load capacity min. 40 kg

## Pre-installation components

### 3.4.1	Ceiling fixture set anchor / bolt mounting for ceiling supply unit

- adaptable on site to height of false ceiling,

- containing all necessary fixation components, e.g. sub-ceiling flange,

- distance tubes, reinforcement set, inter-connecting flange

22  - Ceiling Supply Unit Anaesthesia side with Anaesthesia machine mount.

# 1.1 General Standard of Execution, Regulations

The corresponding requirements of the European Norm EN 793 (special requirements for the safety of medical supply units) Medical Device Directive (MDD) 93/42 EEC as well as the relevant rules and regulations of the law for technical work equipment (equipment safety law) and

ISO/IEC (see following table) in their valid version must be fulfilled through the following specified equipment.

Qualified vendor must have a Quality Management System acc. to ISO 9001 and ISO 13485 (Medical) in place, as well as a MDD (Medical Device Directive) annex 2 section 3 certification

  - Individual Standards and Regulations

| Standards | Basic requirement of the Medical Device Directive (MDD 93/ 42/EWG) • DIN EN 60601-1 / medical electrical devices • DIN EN 60601-1-2 / medical electrical devices • DIN EN 793 / medical supply units • EN ISO 14071 / medical products – Risk analysis |
| --- | --- |
| Classification | Class II b acc. MD Directive 93/42/EWG appendix IX UMDNS Code: 18-046 |

- The equipment must be CE-certified or has got the approval of the European regulatory authority according to MDD.

- The quotation must include all facilities and equipment components, including necessary accessories which are required for the functionality of the equipment.

- Variations or alternatives to the above-mentioned standards or to the following specified scope of performance must be specifically marked with the abbreviations  ”Var.” or ”Alt”.

# Scope of Medical Supply Unit Equipment

The supply unit must meet the high standards for a safe and reliable operation in the OR .

The supply unit must have upgrading capability of new functions for adaptation to future requirements.

- All surfaces resistant against corrosion and disinfectants.

- Colour of painted surfaces: RAL 9002

The functional units of the supply unit are to be completely assembled and tested by the manufacturer.

## General Requirements for Ceiling Supply Units

Ergonomic, optimized supply unit consisting of the following functional units:

- Ceiling fixture set for installation to concrete ceiling structure.

- Ceiling interface for connection of supply unit with electrics and gases delivered from site.

- Horizontal swiveling and electrically-driven height-adjustable arm system.

- Media column with terminal units for medical gas supply/gas evacuation,
high- and low voltage power supply according to specification.

- Double braking system, consisting of friction brake and pneumatic brake for fixation of adjusted arm position.

- The previously mentioned functional units must be coordinated mechanically, functionally and ergonomically, thus presenting a complete medical supply unit.

Scope of Delivery- and Technical Performance

## Basic Units

## 2.1.1	Double Ceiling Supply Unit

- Supply unit with two arms

- Swivel range of arms 320° - 330° at each joint

- End stop adjustable at site according to customer requirements

- Easy arm positioning in horizontal level, also with maximum load

- Color RAL 9002

## Arm Systems

### 2.3.1	Arm System with electromotoric height adjustment;
Arm Length 500+1000 
with Media Column  1000

- Arm system with one horizontal swivel arm, one swivel arm with 	
electromotoric height adjustment, intermediate bearing and ceiling bearing

- Double braking system, consisting of friction brake and pneumatic brake
In case of pressure drop / pressure failure, the arm system remains in it’s previously adjusted position

- Action radius 1500mm

- Lifting time from lowest to highest position max. 20 sec

- Load capacity on media column min. 180 kg

Power supply from site for electrically-driven lift motor

- 230 V

- 1-phase

## Equipment Carrier Systems

### Media Column

- Height of media column  1000

- Fitting of installation components possible on 4 sides

- Pivot bearing integrated into media column

- Swivel range 320 - 330 °

- Adaptation to ceiling height via distance tube attached to arm system

- Mounting system for attachment of workplace components, 	
e.g. shelves, drawers, side arms, equipment holders
from front side

- Working height of workplace components infinitely variable

- Set control elements, consisting of 2 handles for comfortable operation of the supply unit

- Ergonomically designed handles to support the user positioning the system

- Control of all arm functions (e.g. height adjustment, pneumatic brakes) via integrated push buttons

- Color coding of control elements adapted to other supply unit components (e.g. joints, keypads) allowing for intuitive operation

- Mounting at any front plate of media column possible

- One-piece front plates without gaps

- Surfaces resistant against disinfectants

- Color of media column: anodized aluminium

- Color of front plates: anodized aluminium

## Installation Components

### 2.5.1	Gas Fittings

Media column equipped with:

2 x  O2	(oxygen)

1 x Air 5	(compressed air 5 bar)

1 x VAC	(vacuum)

1 x N2O	(nitrous oxide)

1 x AGSS	(Anesthesia Gas Scavenging System)

### 2.5.2	High voltage electrical installation

Media head / media column equipped with:

6 x electrical socket

### 2.5.3	Low voltage electrical installation

Media column equipped with:

1 x 1.	

1 x1

## Workstation Components

2.6.1	CVP-Pole

Support tube with height adjustable extension pole

Length of support tube min. 1000 mm

Length of extension pole min. 800 mm

4 hooks, cross orientation, for accommodation of infusion bottles

2.6.2	Extension arm for mounting at front rail of Media Column

for attachment of a small equipment pole Ø25 mm or CVP-pole

with locking screw on each swivel joint

unrestricted horizontal movement around support tube axis

Length min. 300 mm

Load capacity min. 40 kg

2.6.3	Shelf for mounting at front rail system of Media Column

Shelf for attachment to front rail system of previously described media column  - 3 nos.
Surface area 	440 x 480 mm (W x D)		
Load capacity min. 40 kg	
Thickness max. 22 mm

Material: Aluminium

Smooth, scratchproof, fade-resistant and antistatic surface.

No sharp edges and grooves.

Integrated rubber elements at all 4 edges for injury prevention.

With curved, ergonomic design.

Sida rails should be provided with each shelf.

Drawer for mounting at shelf – 1 Pcs.

2.6.4	Drawer for attachment to shelf, as previously described
Dimensions 	440 x 340 x 170 mm (W x D x H)

Load capacity min. 5 kg

Special drawer runner system with

exceptionally smooth sliding function

self-closing function

Facilitated cleaning by

easy removal capability of complete drawer

removable front fascia

smooth surfaces without protruding handles

Future exchange of front fascia possible due to retro-fit capability

## Pre-installation components

### 2.7.1	Ceiling fixture set anchor / bolt mounting
for ceiling supply unit

- adaptable on site to height of false ceiling,

- containing all necessary fixation components, e.g. sub-ceiling flange,

- distance tubes, reinforcement set, inter-connecting flange

23. Specification of Super-speciality Advance Anaesthesia Workstation

- Should be three gas integrated anaesthesia workstation for ultra modern and super speciality surgical workplaces.

- Should offer   ICU quality ventilator, single user interface of 12.1” colour display and electronic gas dosing suitable for adult, children & neonates.

- Single user interface should control and display all parameter including control of modes, display of cylinder pressures etc.

- The machine should be suitable for low and minimal flow anesthesia application

- The machine should have automatic calculations and presetting of patient specific ventilation settings via ideal body weight.

- The machine should have fully automatic menu drive self test and user check list. In case of emergency it should be possible to bypass the self test completely. The leak test should be done till the Y piece and should not require user interaction during the test.

- It should have configurable screen layouts for individual screen set ups.

- It should have Nitrous oxide free operation configurable.

- The anaesthesia machine, vaporizer and patient monitor should be manufactured by same company.

- System should confirm to EN 60601-2-13 (Requirement for safety and essential performance of anaesthesia system)

- Gas delivering system :

- Unit should have primary connection for Central gas supply for Air, O2 & N2O and as a backup, machine should also have provision for connecting oxygen and nitrous oxide pin index cylinders.

- Machine should have Electronic fresh Gas delivery & monitoring system with automatic switchover to air in case of failure of Oxygen.

- Audio - visual alarm for failure of Oxygen.

- Control of minimum 25% Oxygen in fresh gas upto Flow > 1L/min. and atleast 250 ml of Oxygen concentration for minimal flow application (fresh gas flow < 1L/min). No basal flow.

- Integrated Oxygen flush with self returning valve.

- O2 safety flow adjustable from 0 to 12 L/min, running through the Vaporiser for emergency backup use

- Machine should also have an independent fresh gas outlet for connection to Bain’s or Magill circuit.

- Fresh gas flow setting from 50ml/min to 12 Ltr/min

- It should have the indicator to show the efficiency of fresh gas setting while used in Low flow and minimal flow

- The machine should have Nitrous oxide free operation configurable.

- Breathing system :

- Compact breathing system suitable for minimal flow anaesthesia, with least   patient circuit volume including absorber etc. approx. 3 L (excluding bag) for fast response to change in fresh gas composition.

- Fresh gas decoupled / compensated breathing system for adult and children

- Integrated warmer for breathing gas conditioning and avoidance of condensation.

- APL valve with direct setting of release pressure.

- Integrated Ventilator :

- Electronically controlled electrically driven ventilator should not require any driving gas

- Ventilator suitable for adult and children without changing of bellow.

- Automatic breathing circuit Compliance correction. Light weight bellow should not offer any constant PEEP.

- Spont. Breathing

- Manual Ventilation

- IPPV with Plateau adjustment from 0 to 60% of Ti

- PLV with decelerating flow

- Pressure controlled ventilation

- SIMV in VCV & PCV

- Pressure Support, PS with CPAP, PS with SIMV in VCV/PCV

- Autoflow – delivering set tidal volume at minimum airway pressure.

- High peak inspiratory flow upto 150 LPM,

- Tidal volume adjustment range 5 ml to 1400 ml

- PEEP from 0 to 20 mbar Electronically adjustable

- Resp frequency from 3 to 100 per min.

- I:E : max 5:1 to 1:99

- Should be able to ventilate with atmospheric air, in case of Oxygen and power failure

- Ventilator monitoring :

- Monitoring of Volume, Pressure & Oxygen

- Should monitor patient compliance

- Should curves & Loops with fast analysis of change in lung mechanics provided by simultaneous display of reference & real time loops

- Should display EtCO2, O2 as well as anaesthetic agent and automatic identification of agent

- Bar graph display of volumeter, tidal volume, and the difference of inspiratory and expiratory tidal volume

- Trending for alarms and measured  values with graphic trend display

- The machine should calculate agent consumption and agent uptake by patient on a case by case basis and display of fresh gas consumption in the unit logbook

- Vaporizers :

- Temperature / pressure compensated and flow independent vaporizer. Vaporizer should have extended delivery range from 0 to 6 Vol. %. Vaporizer should have transport lock to provide hermetic sealing of agent chamber during transport & storage.

- The vaporiser design should be maintenance free.

- Alarms:

- The machine should have adjustable alarm limits for all the parameters with auto set alarm function.

- The machine should have automatic display of MAC/MACx values and automatic activation of low agent alarm

- Airway pressure low limit should automatically changes with the PEEP setting changes

- Machine should have 3nos. RS 232 connectivity port for interface to patient monitor / HIS for automatic data acquisition.

- Scope of Supply :

- 3 gas Anaesthesia workstation

- Trolley with drawer

- Writing surface

- Pin Index yokes for O2 & N2O

- Pipeline connections for all three gases

- Integrated Ventilator & monitor

- Semiclosed breathing system

- Adult & Peadiatric autoclavable patient tubings

- Anaesthetic mask size – Adult & child

- Vaporisers for Isoflourane & Sevoflourane

- Central gas supply hoses (Color coded)

- Instruction for use